**Biscayne Bay Aquatic Preserves (BBAP)** **Water Quality Metadata**

**January 2023– September 2023**

**Latest Update:** September 30th, 2023

Note: This is a provisional metadata document; it has not been authenticated as of its download date. Contents of this document are subject to change throughout the QAQC process and it should not be considered a final record of data documentation until that process is complete. Contact the Aquatic Preserve office general email at Biscayne.Bay@FloridaDEP.gov or the Water Quality Specialist at Christopher.Mayer@FloridaDEP.gov with any additional questions.

**I. Data Set and Research Descriptors**

**1) Principal investigator and contact persons –**

Griffin Alexander - Biscayne Bay Aquatic Preserves Manager

(786) 798-4511, Griffin.N.Alexander@FloridaDEP.gov

1277 NE 79th St., Miami FL, 33138

* Project supervisor, field assistance, grant report writer

Christopher Mayer - Water Quality Specialist

(786) 798-4512, Christopher.Mayer@FloridaDEP.gov

1277 NE 79th St., Miami FL, 33138

* Datasonde specialist, telemetry coordinator, water quality data analyst, Quality Assurance/ Quality Control (QAQC) officer, data handling, field assistance, datasonde calibration & deployment, metadata report preparer

Tristan Clute – Water Quality Technician

(786)-798-4515, Tristan.Clute@FloridaDEP.gov

1277 NE 79th St, Miami, FL 33138

* Field and lab logistics manager, field assistance, datasonde specialist, datasonde calibration & deployment, data handling, lab maintenance and organization

**2) Entry verification –**

Deployment data are uploaded from the YSI datasonde to a Personal Computer (IBM compatible). Files are exported from EcoWatch in a comma-delimited format (.CDF), EcoWatch Lite in a comma separated file (.CSV), KOR EXO v2.3.10.0 , or KOR v1.1.8.0 Software in a comma separated file (.CSV) and uploaded to the National Estuarine Research Reserve System (NERRS) Centralized Data Management Office (CDMO) Non-System Wide Monitoring Program (SWMP) Data Upload Service where data undergo automated primary QAQC. All pre- and post-deployment data are removed from the file prior to upload.

During primary QAQC, data are flagged if they are missing or out of sensor range. The edited file is then returned to the Office of Resilience and Coastal Protection (RCP) Data Coordinator and/or the Aquatic Preserve office for secondary QAQC where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro. The macro inserts station codes, creates metadata worksheet for flagged data and summary statistics, and graphs the data for review. It allows the user to apply QAQC flags and codes to the data, remove any overlapping deployment data, append files, and export the resulting data file for upload to the Aquatic Preserve (AP) database.

Upload after secondary QAQC results in incorporation into the AP database as provisional plus data, and finally tertiary QAQC by the RCP’s Data Coordinator and assimilation into the AP database as authenticated data. Where deployment overlap occurs between files, the data produced by the newly calibrated datasonde is accepted as being the most accurate. For more information on QAQC flags and codes, see Sections 11 and 12.

Data was handled by Claire Burgett from January 2019 - January 2020 and then by Aliza Karim from February 2020 - May 2022. Data was handled by Ellyn Willse from May 2022 – July 2023. Data was handled by Christopher Mayer from July 2023 to present.

**3) Research objectives –**

This three-step research project was initiated in response to a 2,000-3,000-acre seagrass loss event in Northern Biscayne Bay, specifically in and around the Julie Tuttle Basin (Basin). Stations were chosen in an array to understand the relationship between the Little River flow into the Basin and across the large, central shoal containing seagrass habitat. Stations with historical data (not necessarily datasonde data) sampled by other agencies were chosen for this project, when appropriate. Not all parameters are taken at all sites across the project, however some stations included sampling for all three research steps. Four sonde stations are set up down the Miami River and across the Rickenbacker Basin to monitor flow from the Miami River across the basin.

* Currently, seven deployed datasonde stations monitor abiotic parameters at 15-minute intervals.
* In addition to the datasondes, this project involves taking monthly water quality grab samples at 23 sites with both surface and bottom samples taken at two sites, LR03, and MRSC. There are 41 parameters collected during water quality grab sampling, including nutrients, pharmaceuticals, herbicides, fungicides, insecticides, and chlorophyll-a. All datasonde stations are also monthly water quality grab sampling stations.
* The 34 benthic assessment sites include seagrass and macroalgae tissue sampling for elemental analyses and stable isotopes, sediment sampling & depth, Braun-Blanquet & cover abundance surveys, and some abiotic factors. All 34 sites are sampled semiannually.

**4) Research methods –**

YSI datasondes are deployed monthly at eight locations, BBLR03, BBBB14, BBJT71, BBMRSC, BBMRDW, BBMRRB, and BBCWA4. Data is collected in 15-minute intervals. All sondes are calibrated before deployment and a pre-deployment calibration verification (ICV) is done as a QAQC check. Approximately every month, the deployed sondes are removed from the water and returned to the lab for data retrieval, cleaning, verifying post-deployment calibration (CCV) and conducting any necessary maintenance or repairs. Freshly calibrated sondes are swapped at the time of retrieval.

All time is reported as Eastern Standard Time.

YSI 6600 and 6920 datasondes have been phased out but were operating in the Bay until April 15th, 2020. Currently, YSI EXO2 (10m depth) sondes are swapped between each station. The sonde is placed into a horizontally deployed PVC pipe with holes to allow for water flow. The unit is mounted to a concrete slab or dock piling and submerged.

BBAP was granted EPA funding to expand the water quality program into the Miami River and Rickenbacker basin in 2019. The seagrass in this basin showed resilience to the die-off event record in 2019 but there is more stakeholder/researcher concerns due to the outflow of water through the Government Cut inlet and onto the coral reef tracks. With this EPA funding, we established two sites in the Miami River in March 2021 and two sites in the Rickenbacker basin in January 2022.

In the Miami River, downstream of a tributary that is known to release nutrient/metal pollutants, an EXO2 site was established on March 31st, 2021 (site: BBMRDW). The mount is attached to a dock and held at depth via an anchor where the sonde is retrieved by lifting the PVC mount out of the water. At the salinity control structure of the Miami River, which is about 9 miles upriver, an EXO2 site was established on March 26th, 2021 (site: BBMRSC). The mount is attached to a concrete seawall and the sonde is retrieved by lifting the PVC mount out of the water.

Near the mouth of the Miami River, EXO2 site BBMRRB was established January 13th, 2022. The sonde is placed in a vertically deployed mount that is constructed from PVC (with flow through holes) attached to an FWC owned piling with custom brackets. EXO2 Site BBCWA4 is located on the edge of the Bill Sadowski Critical Wildlife Area in the same basin as BBMRRB and was established January 18th, 2022. The sonde is placed in a vertically deployed mount that is constructed from PVC (similar to the BBMRRB site) The locations for sites BBMRSC, BBMRDW, BBMRBB and BBCWA4 follow the path of water down the Miami River and out into the Bay.

Specific conductance, pH, turbidity, depth, chlorophyll and dissolved oxygen are calibrated on the EXO2s while temperature is calibrated in factory settings. Calibrations are conducted in the BBAP lab by BBAP staff within 24hours of deployment.

Specific conductance is 1-point calibrated with the 50,000 uS/cm standard (from Ricca). The ICV is done with the 1,000 uS/cm standard while the CCV is done with the 100,000 uS/cm standard. pH is 2-point calibrated with the pH7 standard and pH10 standard (from Fisher Scientific). The ICV and CCV are done with one of the 2 standards. Turbidity is 2-point calibrated with 0 FNU/NTU Milli-Q water and 124 FNU standard (from YSI). The ICV and CCV are done with both the 0 and the 124/126 FNU/NTU standard. Chlorophyll is 2-point calibrated with 0 ug/L Milli-Q water and a rhodamine WT dye standard (from Kingscote). The Tal-PE sensor on the EXO 2 is calibrated for chlorophyll-a (ug/L and RFU). Chlorophyll calibrations using rhodamine WT dye allow for consistency in values across our sondes and deployments. During monthly water quality sampling, sonde readings are paired with water bottle grabs that are tested for chlorophyll using a fluorometric method, allowing for verification of sonde chlorophyll readings.

**5) Site location and character –**

All sites are in relatively close proximity and part of the same watershed. The watershed is an urban, primarily residential part of Miami-Dade County. There are septic tanks in the watershed draining into this area and repeated issues with sewage overflows or other spills. Tidal Range is around 2 feet.

BBLR03 is in between the mouth of the Little River and an upriver salinity control structure and is therefore tidally influenced. The Little River is suspected to be a major source of nutrients and other pollutants into the Basin and generally has poorer water quality than the bay sites. This site was part of Miami-Dade County’s long-term water quality monitoring BayRun program but monitoring at this location was suspended. Freshwater inflow is determined by the salinity control structure and varies depending on management actions, rainfall, and timed releases of Lake Okeechobee water from the salinity control structures.

* Location: 25.846841, -80.182861
* Salinity range: 0.0 – 35.0 ppt
* Depth: 0.01 – 1.48 m (it’s about 10 feet in center of river; sonde is up against the seawall)
* Bottom habitat: bare muddy sediment
* Pollutants detected in 2023:
	+ Surface:
		- **Pesticides/Herbicides:** 2,4 D, Bentazon, Diuron, Fluridone, Imazapyr, Afidopyropen, Imidacloprid
		- **Pharmaceuticals:** Acetaminophen, Carbamazepine, Naproxen
		- **Artificial Sweeteners:** Sucralose, Acesulfame K
	+ Bottom
		- **Pesticides/Herbicides**: 2,4 D, Bentazon, Diuron, Fenuron, Imazapyr
		- **Pharmaceuticals:** Naproxen, Carbamazepine
		- **Artificial Sweeteners:** Sucralose

BBBB14 is in the Northern Biscayne Bay Basin in the northwest section of the shoal. This site was highly impacted by the seagrass die-off event in the northern portion of the shoal where full loss of the prior seagrass cover was observed. The seagrass loss event occurred as an edge-in effect, especially from the northern edge. BBBB14 was likely one of the first sites to switch to a macroalgal dominated bed. The site is part of Miami-Dade County’s long-term water quality monitoring BayRun program. Benthic sampling by BBAP at this site began in 2019 and involves cardinal direction quadrat throws. It’s a marine site with freshwater inflow primarily coming from rainfall and secondarily from the nearby Little River but also has exchange with other basins in Biscayne Bay influenced by other rivers and canals. Seagrass has never been observed in quadrats during our 2019 benthic surveys at this site, however extremely sparse *Syringodium filiforme* was seen in March 2020, while visiting the site. Historical benthic data was not taken at this station, however sites nearby showed high seagrass cover prior to the die-off event.

* Location: 25.83003, -80.15860
* Salinity range: 14.6 – 36.0 ppt
* Depth: 0.12 – 2.0 m
* Bottom habitat: Currently dominated by *Halimeda discoidea* (previously seagrass habitat). Seagrass began reappearing at the site in early 2020 and there are currently patches of *Syringodium filiforme*. Sediment is largely made up of Halimeda hash (remains of calcified algal discs) at the surface in a muddy matrix.
* Pollutants detected in 2023:
	+ Bottom:
		- **Pesticides/Herbicides:** 2,4 D, Bentazon**,** Diuron**,** Imazapyr
		- **Pharmaceuticals:** Carbemazepine
		- **Artificial Sweeteners:** Sucralose

BBJT71 is in the Northern Biscayne Bay Basin in the southeast section of the shoal, further from the Little River than BBBB14. This site was chosen to observe the area where some seagrass still survived or potentially had not yet been impacted by the die-off event at the initiation of the project. The southeast interior section of the shoal in the Basin was previously fully covered by dense refugia of seagrasses, with the dominant species being *Syringodium filiforme*. While there is some seagrass remaining at this site, it is very patchy and sparse. This site was impacted later than BBBB14. Historical data showed averages of 50-75% *Syringodium filiforme* cover before the die-off event, but now has on average 1% *Syringodium filiforme* cover and 3% *Halodule wrightii* cover. Historical data at this site is primarily composed of benthic surveys, but there was a short term datasonde deployment by county researchers. This site was sampled originally by United States Geological Survey (USGS**)** during Fish and Invertebrate Assessment Network surveys in 2014. Miami-Dade County then resampled this location following the seagrass die-off event. Our program established a transect in 2018 at this location for quarterly benthic surveys.

* Location: 25.821730, -80.151250
* Salinity range: 24.2 – 35.9 ppt
* Depth: 0.42 – 2.24 m
* Bottom habitat: Currently dominated by *Halimeda discoidea* but has a mix of seagrasses and other macroalgal species. Sediment is muddy sand with some Halimeda hash (remains of calcified algal discs).
* Pollutants detected in 2023:
	+ Bottom:
		- **Pesticides:** Bentazon**,** Imazapyr, 2,4-D, Diuron
		- **Pharmaceuticals:** Carbemazepine
		- **Artificial Sweeteners:** Sucralose

BBMRDW is located about 1.5-miles upriver from the mouth of the Miami River, downriver of the intersection with Wagner Creek, which is suspected to be a major source of nutrients and pollutants (such as heavy metals). The Miami River outflows into the Rickenbacker Basin shoal and then out the Government Cut inlet which leads out to an offshore coral reef tract. The site is tidally influenced in the wet season and well mixed in the dry season. Water quality in the area is poor with nutrient sources coming from urban run-off and development. There is also significant amount of cargo ship traffic through the Miami River.

* Location: 25.7773, -80.2040
* Salinity range: 0.7 – 34.2 ppt
* Depth: 0.35 – 1.92 m
* Bottom habitat: bare muddy sediment with some hard substrate lining the seawall
* Pollutants detected in 2023
	+ Bottom
		- **Pesticides/Herbicides**: 2,4 D, Bentazon, Diuron, Fluridone, Imazapyr, Imidacloprid
		- **Pharmaceuticals:** Acetaminophen, Carbamazepine
		- **Artificial Sweeteners:** Sucralose, Acesulfame K

BBMRSC is located about 9.1-miles upriver from the mouth of the Miami River, downriver of a salinity control structure. The potential nutrient and pollutant sources are from urban run-off. The Miami River outflows into the Rickenbacker Basin shoal and then out the Government Cut inlet which leads out to an offshore coral reef tract. The site is tidally influenced in the wet season and well mixed in the dry season. Freshwater inflow from the salinity control structure varies depending on management actions, rainfall, and timed releases of Lake Okeechobee water.

* Location: 25.8072, -80.2599
* Salinity range: 0.2 – 26.2 ppt
* Depth: 0.89-2.48 m
* Bottom habitat: bare muddy sediment
* Pollutants detected in 2023
	+ Surface:
		- **Pesticides/Herbicides:** Bentazon, Diuron, Fluridone, Imazapyr, 2,4-D, Imidacloprid
		- **Pharmaceuticals:** Acetaminophen, Carbamazepine, MCPP
		- **Artificial Sweeteners:** Acesulfame-K, Sucralose
	+ Bottom
		- **Pesticides/Herbicides**: Bentazon, Diuron, Fluridone, Imazapyr, 2,4-D, Imidacloprid
		- **Pharmaceuticals:** Acetaminophen, Carbamazepine
		- **Artificial Sweeteners:** Sucralose

BBCWA4 is located in the Rickenbacker Basin shoal on the western edge of the Bill Sadowski Critical Wildlife Area. This site is tidally influenced and receives water from offshore via Government cut inlet and freshwater input from the Miami River. This site is adjacent to an area of dense seagrass habitat.

* Location: 25.758020, -80.167690
* Salinity range: 27.6-37.6
* Depth: 3.07-1.93
* Bottom habitat: mixed seagrass and sponge habitat, near dense seagrass
* Pollutants detected in 2023
	+ Bottom
		- **Pesticides/Herbicides**: Imazapyr, 2,4-D, Bentazon
		- **Pharmaceuticals:**
		- **Artificial Sweeteners:** Sucralose

BBMRRB is located in the Rickenbacker Basin at the confluence of the Miami River and the bay. This area receives heavy boat traffic. This site was selected to monitor water flowing out of the Miami River as it makes its way into the basin and offshore via Government cut inlet.

* Location: 25.769800, -80.180850
* Salinity range: 25.9-37.3 ppt.
* Depth: 2.65-1.40 m
* Bottom habitat: very sparse seagrass, mostly bare ground
* Pollutants detected in 2023
	+ Bottom
		- **Pesticides/Herbicides**: 2,4 D, Imazapyr, Bentazon
		- **Pharmaceuticals:**
		- **Artificial Sweeteners:** Sucralose

BBAP Station Deployment Timeline:

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Station Code | Station Name | Location | Active Dates | Notes |
| BBLR03 | Little River (LR03) | 25.846841, -80.182861 | 4/19/2019 - present | Model 6920 (Jan 1-Mar 17, 2020),EXO2 (Mar 17, 2020-present) |
| BBBB14 | Tuttle Basin NW (BB14) | 25.830030, -80.158600 | 1/18/2019 - present | Model 6600V2-4 (Jan 1-Apr 15, 2020),EXO2 (Apr 15, 2020-present) |
| BBJT71 | Tuttle Basin SE (JT71) | 25.821730, -80.151250 | 3/5/2019 - present | Model EXO2 |
| BBMRDW | Miami River (MRDW) | 25.7773, -80.204 | 3/31/2021 – present | Model EXO2 |
| BBMRSC | Miami River (MRSC) | 25.8072, -80.2599 | 3/25/2021 - present | Model EXO2 |
| BBMRRB | Rickenbacker Basin (MRBB) | 25.76980, -80.18085 | 1/13/2022 - present | Model EXO2 |
| BBCWA4 | Rickenbacker Basin (CWA4) | 25.758020,-80.167690 | 1/18/2022 - present | Model EXO2 |

**6) Data collection period –**

**BBBB14**

|  |  |  |
| --- | --- | --- |
| **Deployment date/time** | **Retrieval date/time** | **Notes** |

 12/7/2023 9:45 1/3/2024 11:00

 1/3/2024 12:00 1/30/2024 9:15

 1/30/2024 9:45 2/26/2024 9:00

 2/26/2024 9:30 3/21/2024 7:45 Sonde dead upon retrieval, battery compartment compromised.

 3/21/2024 8:15 4/15/2024 9:30

 4/15/2024 10:00 5/9/2024 7:00

 5/9/2024 7:30 6/5/2024 8:00

 6/5/2024 8:30 7/3/2024 8:45

 7/3/2024 9:15 8/1/2024 9:30

 8/1/2024 11:00 8/27/2024 8:45

 8/27/2024 10:00 9/24/2024 10:30

 9/24/2024 10:45

**BBJT71**

|  |  |  |
| --- | --- | --- |
| **Deployment date/time** | **Retrieval date/time** | **Notes** |

 12/7/2023 10:45 1/3/2024 12:30

 1/3/2024 13:00 1/30/2024 10:00

 1/30/2024 10:30 2/26/2024 10:00

 2/26/2024 10:15 3/21/2024 9:30

 3/21/2024 9:00 4/15/2024 8:45

 4/15/2024 9:00 5/9/2024 8:15

 5/9/2024 8:30 6/5/2024 8:30

 6/5/2024 9:00 7/3/2024 9:45

 7/3/2024 10:15 8/1/2024 8:30

 8/1/2024 9:00 8/27/2024 9:30

 8/27/2024 10:00 9/24/2024 9:30

 9/24/2024 10:00

**BBLR03**

|  |  |  |
| --- | --- | --- |
| **Deployment date/time** | **Retrieval date/time** | **Notes** |

 12/18/2023 10:00 1/15/2024 8:30

 1/15/2024 9:15 2/8/2024 9:00

 2/8/2024 9:30 2/29/2024 8:45

 2/29/2024 9:15 3/28/2024 8:00

 3/28/2024 8:30 4/17/2024 7:30

 4/17/2024 8:00 5/16/2024 12:30

 5/16/2024 13:00 6/25/2024 8:30

 6/25/2024 9:15 7/18/2024 8:00

 7/18/2024 8:30 8/15/2024 7:45

 8/15/2024 8:00 9/10/2024 8:30

 9/10/2024 8:45

**BBMRDW**

|  |  |  |
| --- | --- | --- |
| **Deployment date/time** | **Retrieval date/time** | **Notes** |

 12/12/2023 8:45 1/5/2024 9:00

 1/5/2024 9:30 1/24/2024 8:45 Replaced DCP SOA

 1/24/2024 9:15 2/21/2024 8:45

 2/21/2024 9:45 3/15/2024 8:15 WB on but sensor faces dirty. Batteries corroded but sonde was alive.

 3/15/2024 9:00 4/9/2024 8:00

 4/9/2024 8:45 5/7/2024 8:15

 5/7/2024 9:00 5/29/2024 7:15

 5/29/2024 8:00 6/27/2024 7:30

 6/27/2024 8:00 7/24/2024 7:45

 7/24/2024 9:00 8/22/2024 8:45

 8/22/2024 9:30 9/19/2024 10:00

 9/19/2024 10:30

**BBMRSC**

|  |  |  |
| --- | --- | --- |
| **Deployment date/time** | **Retrieval date/time** | **Notes** |

 12/12/2023 8:45 1/5/2024 10:00

 1/5/2024 10:45 1/24/2024 9:45 Cleaned chain and reinforced with zip ties

 1/24/2024 10:15 2/21/2024 10:15 Faces dirty, wiper brush off

 2/21/2024 10:45 3/15/2024 9:30 Telemetry failed. Sonde faces dirty, sonde completely dead

 3/15/2024 10:00 4/9/2024 9:15

 4/9/2024 9:15 5/7/2024 9:30

 5/7/2024 10:00 5/29/2024 8:30

 5/29/2024 9:00 6/27/2024 8:45

 6/27/2024 9:15 7/24/2024 10:15

 7/24/2024 11:15 8/22/2024 6:45

 8/22/2024 7:30 9/19/2024 11:00

 9/19/2024 11:15

**BBMRRB**

|  |  |  |
| --- | --- | --- |
| **Deployment date/time** | **Retrieval date/time** | **Notes** |

 12/22/2023 10:15 1/19/2024 9:00

 1/19/2024 9:45 2/13/2024 9:45

 2/13/2024 10:45 3/7/2024 11:30 Amphipods on CT sensor and silt on TB sensor

 3/7/2024 12:00 4/4/2024 9:15 Uploaded new program to datalogger, connected manually.

 4/4/2024 10:00 5/2/2024 8:45

 5/2/2024 9:30 5/24/2024 7:45

 5/24/2024 10:00 6/21/2024 9:15

 6/21/2024 9:45 7/11/2024 11:15

 7/11/2024 11:45 8/8/2024 8:15

 8/8/2024 9:00 9/6/2024 8:45

 9/6/2024 9:15

**BBCWA4**

|  |  |  |
| --- | --- | --- |
| **Deployment date/time** | **Retrieval date/time** | **Notes** |
| 12/8/2022 12:45 | 1/2/2023 12:45 |  |
| 1/2/2023 13:15 | 1/30/2023 11:15 |  |

 12/22/2023 10:15 1/19/2024 10:00

 1/19/2024 10:30 2/13/2024 11:00

 2/13/2024 12:00 3/7/2024 9:45

 3/7/2024 10:30 4/4/2024 8:30

 4/4/2024 9:00 5/2/2024 7:45

 5/2/2024 8:30 5/24/2024 10:15

 5/24/2024 11:15 6/21/2024 8:15

 6/21/2024 8:45 7/11/2024 10:15

 7/11/2024 10:45 8/8/2024 9:30

 8/8/2024 10:15 9/6/2024 9:30

 9/6/2024 9:45

**7) Distribution –**

The Principle Investigator (PI) retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the Florida Department of Environmental Protection and the Aquatic Preserve site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.

Aquatic Preserve water quality data and metadata can be obtained from the Manager at the individual Aquatic Preserve site (please see Principal Investigators and Contact Persons) and online at the Aquatic Preserve’s data portal home page [www.floridaapdata.org](http://www.floridaapdata.org). Data are available in comma delimited format.

**8) Associated researchers and projects –**

There are several agencies conducting water quality monitoring within Biscayne Bay. The following are the projects that we are aware of though there may be more.

(This Project) Florida Department of Environmental Protection (DEP) – Biscayne Bay Aquatic Preserves (BBAP)

* Currently, seven deployed datasonde stations monitor abiotic parameters at 15-minute intervals.
* In addition to the datasondes, this project involves taking monthly water quality grab samples at 21 sites. The 31 parameters are collected during water quality grab sampling, include nutrients, pharmaceuticals, herbicides, fungicides, insecticides, and chlorophyll-a. All datasonde stations are also monthly water quality grab sampling stations. Data can be acquired from Ellyn Willse.
* The 34 benthic assessment sites include seagrass and macroalgae tissue sampling for elemental analyses and stable isotopes, sediment sampling & depth, Braun-Blanquet & percent cover abundance surveys, and some abiotic factors. Previously quarterly sampling of eleven benthic sites and all twenty-one sites semiannually occurred. Currently there are 34 sites which are monitored semi-annually.
	+ In September 2020, an additional 10 benthic sites were added in the Rickenbacker Basin (the outflow basin of the Miami River). These sites will be sampled for all the above parameters, semi-annually.
	+ In March 2023, benthic monitoring sites were updated. Several sites were added to the JFK Basin, one site was moved and one removed due to boat traffic danger and site RB1A was removed due to distance from other sites. The total number of benthic sites is now 34.
	+ Sediment sulfide analyses have also been done in conjunction with the SAV sampling at all sites in September 2020.
	+ In April 2021, grain size analyses were conducted in the Rickenbacker basin and sedimentary iron sampling was done at all benthic sites in both basins.
* <https://floridadep.gov/rcp/aquatic-preserve/locations/biscayne-bay-aquatic-preserves>
* This project has expanded to the Rickenbacker basin and Miami River under an EPA grant. Four EXO2 datasonde sites were added: 1 in the basin, 1 at the intersection of the river mouth and boating channel and 2 upriver.
* BBAP has conducted (from August 2020 to present) monthly bird surveys on 4 islands within Northern Biscayne Bay in response to a bird rookery collapse on Bird Key island. Data can be acquired from Griffin Alexander.
* Telemetry was added to all sites summer 2022 and live data will be available online through a DEP hosted website in the future.

Florida Department of Environmental Protection (DEP) – Office of Resilience and Coastal Protection (ORCP)

* ORCP is handling the establishment and characterization of the Government Cut Inlet Contributing Area (ICA) that includes the upland watershed that flows into Biscayne Bay, Biscayne Bay itself, and the coastal reef area just offshore of the government cut inlet. The watershed will be characterized by partners at NOAA via increased water quality monitoring in collaboration with this BBAP project, Miami-Dade County’s BayRun project, and NOAA’s pollutant loading model development. Methodology development conversations are ongoing with a planned initiation in fall 2021. The program’s goals include addressing EPA’s 9 elements of a comprehensive Watershed Management Plan for the ICA.
* Further research will be initiated through a DEP Coral Reef Protection and Restoration Grant program that will be awarded in late 2020/early 2021

Florida International University (FIU)

* Datasonde monitoring via the Center for Aquatic Chemistry and Environment (CREST) program with live telemetry data available in sites located in the same basin and adjacent basins as the BBAP sondes available through Dr. Piero Gardinali.
* <https://crestcache.fiu.edu/research/research-buoys/index.html>
* Long-term water quality sampling within Biscayne Bay as part of the Southeast Environmental Research Center (SERC) Water Quality Monitoring Network.
* <http://serc.fiu.edu/wqmnetwork/>
* EPA funded project on hydrodynamic modeling of Biscayne Bay with preliminary results available through Dr. Henry Briceno.
* FIU contract via the seagrass lab through Dr. Jim Fourqurean to assess BBAP collected samples of sediments and seagrass and algae tissues.
* Various graduate student projects within and around Biscayne Bay.

Miami-Dade County Department of Environmental Resources Management (DERM)

* Long-term monitoring for environmental data and water quality parameters within Biscayne Bay, canals, and their tributaries (BayRun program).
* Data available via request and through .pdf reports.
* <http://www.miamidade.gov/environment/surface-water-quality.asp>

South Florida Water Management District (SFWMD)

* Long-term surface water quality monitoring at coastal and canal sites in central and south Biscayne Bay.
* Flow monitoring from canal and river outputs around Biscayne Bay.
* Groundwater, sediment, rainfall and weather monitoring data available as well.
* Data stored and accessible through DBHYDRO database.
* <http://my.sfwmd.gov/dbhydroplsql/show_dbkey_info.main_menu>

Miami Waterkeeper (MWK) and Florida Department of Health (DOH)

* DOH Healthy Beaches program supplemented by bay sampling from Miami Waterkeeper
* Sample monitoring for fecal indicator bacteria.
* <http://www.floridahealth.gov/environmental-health/beach-water-quality/index.html>
* Continuous surface parameter data available via deployed buoys in Northern Biscayne Bay.
* <https://crestcache.fiu.edu/research/research-buoys/index.html>

National Oceanic and Atmospheric Administration (NOAA)

* Long-term water quality sampling.
* Turbidity study done in Northern Biscayne Bay, including a station in the Tuttle Basin that included instrument deployments and bottle grabs.
* <https://pdfs.semanticscholar.org/70c8/c68607f1bf40bc13c331fb2d8e09213f5658.pdf>

University of Florida, Institute of Food and Agricultural Sciences (IFAS) Sea Grant Extension Program

* Biscayne Bay Water Watch Program.
* Community-based volunteer monitoring program to take over SFMWD and DERM sites that lost funding.
* Abiotic, nutrient, and chlorophyll a data collected.
* <https://sfyl.ifas.ufl.edu/miami-dade/natural-resources/biscayne-bay-water-watch-/>

University of Central Florida

* eDNA analysis of surface waters at 8- ecologically different sites with Biscayne Bay (4 of which will be in Northern Biscayne Bay). Project to be initiated in Aug. 2020 by Dr. Michelle Gaither’s Lab.
* <https://sciences.ucf.edu/biology/gaitherlab/>
* Rapid Evaluation of Estrogenic Endocrine Disrupting Compounds sampling event through surface sampling at 10 sites in Biscayne Bay. Project to be initiated Fall 2020 by Dr. John Fauth’s Lab.

University of Miami

* Substrate and temporal viability of enterococci bacteria by Dr. Helena Solo-Gabriele and students.
* Various graduate student projects within and around Biscayne Bay.
* <https://www.rsmas.miami.edu/research/projects/index.html>

Kent State University

* Imaging and analysis of fluorescent pigment data collected through satellite imagery by Dr. Joseph Ortiz.
* In preliminary stages but intended to be able to determine the spatial extent of seagrass versus macro-algal dominated beds as well as chlorophyll blooms.

**II. Physical Structure Descriptors**

**9) Sensor specifications –**

Currently, we have 13 YSI EXO2 Sondes that are interchangeably swapped monthly at each station. At BBLR03, prior to March 17, 2020, two YSI 6920s were swapped monthly. At BBBB14, prior to April 15, 2020, one YSI 6600 V2-4 was deployed monthly.

**YSI EXO2 (10m depth) data sonde:**

Parameter: Total algae

Units: micrograms/Liter

Sensor Type: optical sensor

Model#: YSI 599103-01

Range: 0 to 400 µg/L chl

Accuracy: Linearity: r2 ≥ 0.999 for Rhodamine WT across full range

Resolution: 0.01 RFU or 0.01 µg/L chl

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: 4-electrode nickel cell

Model#: YSI 599827

Range:0 to 200 mS/cm

Accuracy: 0 to 100: +/- 0.5% of reading or 0.001 mS/cm, w.i.g.; 100 to 200 +/- 1% of reading

Resolution: 0.0001 to 0.01 mS/cm

Parameter: Temperature

Units: Celsius (C)

Sensor Type: 4-electrode nickel cell

Model#: YSI 599827

Range: -5 to 35 ºC, 35 to 50 ºC

Accuracy: -5 to 35 ºC: +/- 0.01 ºC2, 35 to 50 ºC: +/- 0.05 ºC2

Resolution: 0.001 ºC

Parameter: dissolved oxygen, % air saturation

Units: percent air saturation (%)

Sensor Type: optical sensor

Model#: YSI 599100-01

Range: 0 to 500% air saturation

Accuracy: 0 to 200%: +/1 1% of reading or 1% saturation, w.i.g.; 200 to 500% +/- 5% of reading

Resolution: 0.1% air saturation

Parameter: dissolved oxygen, mg/L

Units: mg/L

Sensor Type: optical sensor

Model#: YSI 599100-01

Range: 0 to 50 mg/L

Accuracy: 0 to 20 mg/L: +/- 0.1 mg/L or 1% of reading, w.i.g.; 20 to 50 mg/L: 5% of reading

Resolution: 0.01 mg/L

Parameter: pH

Units: pH units

Sensor Type: unguarded

Model#: YSI 599702

Range: 0 to 14 units

Accuracy: +/- 0.1 pH units within +/- 10ºC of calibration temp; +/- 0.2 pH units for entire temp range

Resolution: 0.01 units

Parameter: Turbidity

Units: FNU (formazin nephelometric units)

Sensor Type: optical sensor

Model#: YSI 599101-01

Range: 0 to 4000 FNU

Accuracy: 0 to 999 FNU: 0.3 FNU or +/- 2% of reading, w.i.g.; 1000 to 4000 FNU: +/- 5% of reading

Resolution: 0 to 999 FNU=0.01 FNU; 1000 to 4000 FNU=0.1 FNU

Parameter: Depth

Units: meters

Sensor Type: integral, non-vented depth sensor

Range: 0 to 10 m

Accuracy: +/- 0.04 FS (+/- 0.004m or +/-0.013ft)

Resolution: 0.001 ft (0.001 m)

**Sensor Disclaimers:**

**Depth Qualifier:**

YSI data sondes can be equipped with either vented or non-vented depth/level sensors.  Readings for both vented and non-vented sensors are automatically compensated for water density change due to variations in temperature and salinity; but for all non-vented depth measurements, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.02 cm for every 1 millibar change in atmospheric pressure and is eliminated for vented sensors because they are vented to the atmosphere throughout the deployment time interval.

Standard calibration protocol calls for all non-vented depth sensors to read 0 meters at a (local) barometric pressure of 1013.25 mb (760 mm/hg).  To achieve this, each site calibrates their depth sensor with a depth offset number, which is calculated using the actual atmospheric pressure at the time of calibration and the equation provided in the Aquatic Preserve calibration sheet or digital calibration log.  This offset procedure standardizes each depth calibration. If accurate atmospheric pressure data are available, non-vented sensor depth measurements can be corrected. The Principal Investigator should be contacted in order to obtain information regarding atmospheric pressure data availability.

**Salinity Units Qualifier:**

The 6600 series sondes report salinity in parts per thousand (ppt) units, the EXO sondes report practical salinity units (psu). These units are essentially the same and for Aquatic Preserve purposes are understood to be equivalent, however psu is considered the more appropriate designation. Moving forward the Aquatic Preserve program will assign psu salinity units for all data regardless of sonde type.

**Turbidity Qualifier:**

The 6600 series sondes report turbidity in nephelometric turbidity units (NTU), the EXO sondes use formazin nephelometric units (FNU). These units are essentially the same but indicate a difference in sensor methodology, for Aquatic Preserve purposes they will be considered equivalent. Moving forward, the Aquatic Preserve program will use FNU/NTU as the designated units for all turbidity data regardless of sonde type. If turbidity units and sensor methodology are of concern, please see the Sensor Specifications portion of the metadata.

**Chlorophyll Fluorescence Disclaimer:**

YSI chlorophyll sensors (6025 or 599102-01) are designed to serve as a proxy for chlorophyll concentrations in the field for monitoring applications and complement traditional lab extraction methods; therefore, there are accuracy limitations associated with the data that are detailed in the YSI manual including interference from other fluorescent species, differences in calibration method, and effects of cell structure, particle size, organism type, temperature, and light on sensor measurements.

**10) Coded variable definitions –**

Sampling station: Station code:

LR03 BBLR03

JT71 BBJT71

BB14 BBBB14

MRSC BBMRSC

MRDW BBMRDW

CWA4 BBCWA4

MRRB BBMRRB

**11) QAQC flag definitions –**

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F\_). During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is missing and above or below sensor range. All remaining data are then flagged 0, passing initial QAQC checks. During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5 Outside High Sensor Range

-4 Outside Low Sensor Range

-3 Data Rejected due to QAQC

-2 Missing Data

-1 Optional SWMP Supported Parameter

 0 Data Passed Initial QAQC Checks

 1 Suspect Data

 2 *Open - reserved for later flag*

 3 Calculated data: non-vented depth/level sensor correction for changes in barometric pressure

 4 Historical Data: Pre-Auto QAQC

 5 Corrected Data

**12) QAQC code definitions** –

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column. There are three (3) different code categories, general, sensor, and comment. General errors document general problems with the deployment or YSI datasonde, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data. Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an \* below) can be applied to the entire record in the F\_Record column.

General Errors

 GIC No instrument deployed due to ice

 GIM Instrument malfunction

 GIT Instrument recording error; recovered telemetry data

 GMC No instrument deployed due to maintenance/calibration

 GNF Deployment tube clogged / no flow

 GOW Out of water event

 GPF Power failure / low battery

 GQR Data rejected due to QA/QC checks

 GSM See metadata

Corrected Depth/Level Data Codes

 GCC Calculated with data that were corrected during QA/QC

 GCM Calculated value could not be determined due to missing data

 GCR Calculated value could not be determined due to rejected data

 GCS Calculated value suspect due to questionable data

 GCU Calculated value could not be determined due to unavailable data

Sensor Errors

 SBO Blocked optic

 SCF Conductivity sensor failure

 SCS Chlorophyll spike

 SDF Depth port frozen

 SDG Suspect due to sensor diagnostics

 SDO DO suspect

 SDP DO membrane puncture

 SIC Incorrect calibration / contaminated standard

 SNV Negative value

 SOW Sensor out of water

 SPC Post calibration out of range

 SQR Data rejected due to QAQC checks

 SSD Sensor drift

 SSM Sensor malfunction

 SSR Sensor removed / not deployed

 STF Catastrophic temperature sensor failure

 STS Turbidity spike

 SWM Wiper malfunction / loss

Comments

 CAB\* Algal bloom

 CAF Acceptable calibration/accuracy error of sensor

 CAP Depth sensor in water, affected by atmospheric pressure

 CBF Biofouling

 CCU Cause unknown

 CDA\* DO hypoxia (<3 mg/L)

 CDB\* Disturbed bottom

 CDF Data appear to fit conditions

 CFK\* Fish kill

 CIP \* Surface ice present at sample station

 CLT\* Low tide

 CMC\* In field maintenance/cleaning

 CMD\* Mud in probe guard

 CND New deployment begins

 CRE\* Significant rain event

 CSM\* See metadata

 CTS Turbidity spike

 CVT\* Possible vandalism/tampering

 CWD\* Data collected at wrong depth

 CWE\* Significant weather event

**13) Post deployment information** –

CCV values in red did not pass the established acceptance criteria for the Florida AP database. Data from the deployments that did not pass CCV were labeled as suspect. AC was ±2% of standard value for specific conductance, turbidity, DO (mg/L) and chlorophyll. AC was ±1% for DO (%). AC was ±0.2 for pH and temperature (°C). CCV values in dark re~~d~~ that are bold, fell beyond 3x the acceptance criteria so the data from the deployment was rejected. Values in italics are standard values. Standard values for DO (mg/L), Chlorophyll (ug/L), BGA (ug/L) and BGA (RFU) are determined by temperature.

Sondes A-M are EXO2s. Sonde A and B were previously called 4 and 5, respectively. Note: SpCond 70 was sent to us in May, therefore, the acceptance criteria in that month is ±5% of 70.0.

|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| **BBLR03**  | **Parameter**  | **Specific Conductivity (mS/cm)**  | **Specific Conductivity (mS/cm)**  | **pH**  | **Turbidity (FNU)**  |   | **Dissolved Oxygen  (%, mg/L)**  |   |   | **Temperature (°C)**  |   | **Depth (m)**  | **Chlorophyll-a (RFU)**  |   |   | **Chlorophyll-a (ug/L)**  |   |   |
| **Sonde**  | **Date/Standard**  | **50**  | **100**  | **7**  | **0**  | **100%**  | **Sonde**  | **0**  | **0**  | **0**  |
| E  | 1/15/2024  |    | 100.787  | 6.98  | -0.23  | 124.61  | 100.9  | 8.70  | *8.661*  | 22.6  | *22.6*  | 0.061  | 0.05  | 15.29  | *15.92*  | 0.00  | 60.97  | *64.0*  |
| D  | 2/8/2024  |    | 101.098  | 6.96  | 0.05  | 123.58  | 102.6  | 9.25  | *9.04*  | 20.2  | *20.2*  | 0.083  | -0.06  | 14.78  | *16.28*  | -0.27  | 59.54  | *65.5*  |
| B  | 2/29/2024  |    | 101.003  | 6.96  | -0.02  | 123.51  | 101.1  | 8.74  | *8.661*  | 22.4  | *22.5*  | 0.092  | -0.08  | 18.18  | *15.74*  | -0.9  | 73.15  | *63.28*  |
| A  | 3/28/2024  |    | 104.277  | 7.04  | 0.08  | 124.34  | 99.3  | 8.31  | *8.418*  | 24.2  | *24.3*  | -0.01  | -0.02  | 15.18  | *15.44*  | -0.11  | 61.32  | *62.18*  |
| J | 4/17/2024 |  | 100.961 | 7.01 | -0.15 | 125.05 | 100.1 | 8.45 | *8.498* | 23.7 | *23.6* | 0.08 | -0.01 | 14.93 | *15.44* | 0.01 | 60.32 | *62.18* |
| G | 5/16/2024 |  | 70.078 | 7.17 | -0.39 | 123.03 | 102.3 | 8.47 | *8.341* | 24.8 | *24.8* | -0.007 | 0.03 | 15.32 | *15.23* | 0.06 | 61.79 | *61.41* |
| D | 6/25/2024 |  | 99.234 | 7.08 | 0.36 | 123.96 | 100.0 | 8.38 | *8.418* | 24.2 | *24.2* | 0.002 | 0.02 | 15.57 | *15.35* | 0.07 | 62.42 | *61.85* |
| G | 7/18/2024 |  | 99.931 | 6.97 | 0.27 | 123.96 | 100.7 | 8.33 | *8.341* | 24.8 | *24.8* | 0.061 | 0.00 | 14.97 | *15.14* | -0.06 | 60.88 | *61.04* |
| L | 8/15/2024 |  | 100.942 | 7.03 | -0.03 | 123.55 | 100.9 | 7.98 | *7.898* | 27.9 | *27.7* | 0.013 | 0.03 | 14.78 | *15.08* | -0.13 | 59.87 | *60.78* |
| A | 9/10/2024 |  | 100.600 | 6.95 | 0.13 | 123.07 | 99.8 | 8.31 | *8.341* | 24.8 | *24.5* | 0.007 | 0.00 | 15.03 | *15.23* | 0.08 | 60.62 | *61.41* |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| **BBBB14**  | **Parameter**  | **Specific Conductivity (mS/cm)**  | **Specific Conductivity (mS/cm)**  | **pH**  | **Turbidity (FNU)**  |   | **Dissolved Oxygen  (%, mg/L)**  |   |   | **Temperature (°C)**  |   | **Depth (m)**  | **Chlorophyll-a (RFU)**  |   |   | **Chlorophyll-a (ug/L)**  |   |   |
| **Sonde**  | **Date/Standard**  | **50**  | **100**  | **7**  | **0**  | **100%**  | **Sonde**  | **0**  | **0**  | **0**  |
| M  | 1/3/2024  |    | 101.450  | 7.04  | 0.21  | 123.41  | 101.1  | 9.24  | *9.276*  | 19.0  | *19.0*  | 0.032  | -0.01  | 17.05  | *16.73*  | -0.02  | 73.47  | *67.32*  |
| B  | 1/30/2024  |    | 101.409  | 7.99  | 0.13  | 123.94  | 100.3  | 9.26  | *9.184*  | 19.1  | *19.1*  | 0.063  | -0.02  | 15.93  | *16.7*  | -0.1  | 61.18  | *67.2*  |
| A  | 2/26/2024  |    | 100.820  | 6.95  | 0.02  | 124.88  | 99.8  | 9.07  | *9.092*  | 19.9  | *19.8*  | 0.103  | -0.01  | 16.43  | *16.52*  | 0.11  | 63.17  | *66.48*  |
| C  | 3/21/2024  |    | N/A  | N/A  | N/A  | N/A  | N/A  | N/A  | *N/A*  | N/A  | *N/A*  | N/A  | N/A  | N/A  | *N/A*  | N/A  | N/A  | *N/A*  |
| G | 4/15/2024 |  | 100.596 | 6.92 | -0.23 | 125.81 | 100.4 | 8.62 | *8.578* | 22.7 | *22.7* | 0.058 | -0.03 | 12.46 | *15.59* | -0.12 | 50.15 | *62.73* |
| D | 5/9/2024 |  | 70.304 | 7.04 | -0.28 | 124.32 | 100.4 | 8.35 | *8.341* | 24.6 | *24.7* | 0.028 | -0.08 | 14.99 | *14.87* | -0.24 | 60.28 | *59.87* |
| C | 6/5/2024 |  | 100.430 | 6.99 | 0.24 | 112.32 | 101.30 | 8.42 | *8.510* | 24.0 | *24.1* | 0.005 | 0.03 | 15.18 | *15.26* | 0.09 | 61.00 | *61.52* |
| A | 7/3/2024 |  | 101.377 | 7.05 | -0.21 | 123.90 | 99.8 | 8.37 | *8.341* | 22.5 | *22.2* | 0.067 | 0.00 | 15.02 | *15.32* | 0.06 | 59.91 | *61.74* |
| F | 8/1/2024 |  | 101.398 | 7.00 | -0.10 | 126.06 | 100.9 | 8.44 | *8.418* | 24.3 | *24.2* | 0.073 | 0.01 | 17.50 | *15.23* | 0.08 | 70.43 | *61.41* |
| D | 8/27/2024 |  | 102.400 | 7.00 | 2.42 | 126.35 | 101.0 | 8.32 | *8.418* | 24.4 | *24.4* | 0.045 | 0.66 | 15.86 | *15.23* | 2.61 | 63.41 | *61.41* |
| E | 9/24/2024 |  | 99.851 | 6.97 | 0.19 | 124.45 | 99.3 | 8.35 | *8.418* | 24.2 | *24.1* | 0.003 | 0.02 | 16.16 | *15.44* | 0.02 | 65.20 | *62.18* |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| **BBJT71**  | **Parameter**  | **Specific Conductivity (mS/cm)**  | **Specific Conductivity (mS/cm)**  | **pH**  | **Turbidity (FNU)**  |   | **Dissolved Oxygen  (%, mg/L)**  |   |   | **Temperature (°C)**  |   | **Depth (m)**  | **Chlorophyll-a (RFU)**  |   |   | **Chlorophyll-a (ug/L)**  |   |   |
| **Sonde**  | **Date/Standard**  | **50**  | **100**  | **7**  | **0**  | **100%**  | **Sonde**  | **0**  | **0**  | **0**  |
| I  | 1/3/2024  |    | 101.672  | 7.07  | 0.01  | 122.44  | 105.7  | 9.77  | *9.276*  | 19.0  | *19.1*  | 0.030  | -0.05  | 17.52  | *16.76*  | -0.22  | 70.51  | *67.44*  |
| B  | 1/30/2024  |    | 101.143  | 6.97  | 0.17  | 123.84  | 99.6  | 9.18  | *9.184*  | 19.1  | *19.1*  | 0.055  | -0.01  | 15.44  | *16.7*  | -0.05  | 61.82  | *67.2*  |
| A  | 2/26/2024  |    | 101.338  | 6.97  | 0.04  | 124.66  | 100.1  | 9.07  | *9.092*  | 20.0  | *19.8*  | 0.114  | -0.05  | 14.67  | *16.49*  | -0.16  | 55.75  | *66.36*  |
| C  | 3/21/2024  |    | 101.586  | 6.95  | -0.16  | 126.03  | 100.5  | 8.58  | *8.578*  | 23.1  | *23.1*  | 0.081  | 0.04  | 18.03  | *15.56*  | -0.24  | 72.38  | *62.62*  |
| E | 4/15/2024 |  | 101.457 | 7.14 | 0.10 | 124.51 | 99.9 | 8.60 | *8.661* | 22.7 | *22.7* | 0.053 | -0.08 | 17.61 | *15.62* | -0.21 | 67.96 | *62.84* |
| K | 5/9/2024 |  | 70.406 | 6.99 | -0.16 | 124.56 | 100.1 | 8.31 | *8.341* | 24.7 | *24.7* | 0.021 | -0.02 | 14.76 | *14.87* | -0.06 | 59.19 | *59.87* |
| L | 6/5/2024 |  | 101.385 | 6.93 | 0.05 | 124.63 | 100.0 | 8.42 | *8.418* | 24.0 | *24.0* | 0.008 | -0.01 | 14.77 | *15.44* | 0.01 | 59.52 | *62.18* |
| J | 7/3/2024 |  | 101.624 | 7.15 | -0.97 | 123.71 | 100.3 | 8.35 | *8.341* | 24.5 | *24.4* | 0.069 | 0.02 | 16.05 | *15.20* | 0.07 | 62.96 | *61.30* |
| H | 8/1/2024 |  | 101.765 | 7.05 | -0.08 | 124.30 | 101.80 | 8.51 | *8.418* | 24.3 | *24.2* | -0.371 | 0.03 | 17.12 | *15.29* | 0.01 | 68.39 | *61.63* |
| K | 8/27/2024 |  | 96.194 | 6.96 | -0.23 | 125.09 | 100.5 | 8.33 | *8.418* | 24.4 | *24.4* | 0.041 | 0.06 | 15.36 | *15.26* | 0.28 | 61.79 | *61.52* |
| G | 9/24/2024 |  | 99.576 | 7.06 | 0.08 | 128.32 | 98.8 | 8.28 | *8.418* | 24.1 | *24.2* | 0.007 | -0.05 | 15.36 | *15.38* | -0.13 | 61.99 | *61.96* |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| **BBMRDW**  | **Parameter**  | **Specific Conductivity (mS/cm)**  | **Specific Conductivity (mS/cm)**  | **pH**  | **Turbidity (FNU)**  | **Dissolved Oxygen  (%, mg/L)**  | **Temperature (°C)**  | **Depth (m)**  | **Chlorophyll-a (RFU)**  | **Chlorophyll-a (ug/L)**  |
| **Sonde**  | **Date/Standard**  | **50**  | **100 mS/cm**  | **7**  | **0**  | **124**  | **100%**  | **DO CCV**  | ***Std***  | **Sonde**  | ***Therm***  | **0**  | **0**  | **Chl CCV**  | ***Std***  | **0**  | **Chl CCV**  | ***Std***  |
| D  | 1/5/2024  |   | 100.761  | 7.02  | 0.19  | 123.91  | 99.3  | 8.70  | *8.743*  | 21.7  | *21.8*  | 0.054  | 0.02  | 16.25  | *16.07*  | -0.03  | 64.21  | *64.63*  |
| I  | 1/24/2024  |   | 100.517  | 7.02  | -0.01  | 120.6  | 101.7  | 8.77  | *8.578*  | 22.9  | *22.8*  | 0.115  | 0.07  | 13.67  | *15.86*  | 0.14  | 55.36  | *63.75*  |
| C  | 2/21/2024  |   | 101.188  | 7.01  | -0.04  | 125.59  | 102.0  | 9.25  | *9.092*  | 20.0  | *20.0*  | 0.097  | 0.03  | 16.56  | *16.46*  | -0.01  | 66.65  | *66.24*  |
| G  | 3/15/2024  |   | 100.910  | 6.99  | -0.07  | 123.43  | 103.3  | 8.96  | *8.578*  | 22.8  | *22.8*  | 0.062  | 0.02  | 15.74  | *15.65*  | 0.01  | 63.27  | *62.95*  |
| D | 4/9/2024 |  | 101.192 | 7.01 | 0.06 | 125.35 | 100.1 | 8.57 | *8.578* | 22.9 | *23.0* | 0.065 | -0.06 | 14.85 | *15.65* | -0.10 | 59.66 | *62.95* |
| B | 5/7/2024 |  | 100.845 | 7.07 | 0.08 | 123.09 | 105.0 | 8.86 | *8.418* | 23.7 | *23.8* | -0.001 | -0.02 | 11.12 | *15.44* | -0.07 | 44.46 | *62.18* |
| A | 5/29/2024 |  | 100.576 | 7.00 | -0.08 | 124.28 | 101.5 | 8.46 | *8.341* | 24.4 | *24.7* | 0.047 | -0.01 | 15.01 | *15.23* | 0.04 | 60.57 | *61.41* |
| J | 6/27/2024 |  | 101.206 | 7.11 | 0.75 | 121.3 | 99.7 | 8.37 | *8.418* | 24.0 | *24.1* | 0.081 | -0.06 | 14.84 | *15.32* | -0.27 | 60.32 | *61.72* |
| D | 7/24/2024 |  | 100.655 | 7.00 | 0.24 | 123.6 | 101.1 | 8.38 | *8.263* | 24.9 | *24.9* | 0.085 | -0.07 | 17.07 | *15.35* | -0.28 | 68.8 | *61.85* |
| G | 8/22/2024 |  | 101.017 | 7.12 | -0.12 | 122.65 | 100.2 | 8.37 | *8.340* | 24.3 | *24.4* | 0.059 | 0.01 | 15.14 | *15.26* | -0.02 | 61.14 | *61.52* |
| L | 9/19/2024 |  | 101.512 | 7.03 | 0.02 | 124.13 | 100.4 | 8.42 | *8.341* | 24.2 | *24.3* | -0.006 | 0.02 | 15.14 | *15.23* | 0.01 | 60.80 | *61.80* |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|   |   |   |   |   |     |   |   |   |   |   |   |   |   |   |   |   |   |   |
|   |   |   |   |   |   |   |   |   |   |   |   |   |   |   |   |   |   |   |
| **BBMRSC**  | **Parameter**  | **Specific Conductivity (mS/cm)**  | **Specific Conductivity (mS/cm)**  | **pH**  | **Turbidity (FNU)**  | **Dissolved Oxygen  (%, mg/L)**  | **Temperature (°C)**  | **Depth (m)**  | **Chlorophyll-a (RFU)**  | **Chlorophyll-a (ug/L)**  |
| **Sonde**  | **Date/Standard**  | **50**  | **100**  | **7**  | **0**  | **124**  | **100%**  | **DO CCV**  | ***Std***  | **Sonde**  | ***Therm***  | **0**  | **0**  | **Chl CCV**  | ***Std***  | **0**  | **Chl CCV**  | ***Std***  |
| K  | 1/5/2024  |   | 100.283  | 7.09  | 0.22  | 124.15  | 98.9  | 8.66  | *8.743*  | 21.8  | *21.8*  | 0.051  | -0.12  | 14.53  | *16.01*  | -0.43  | 58.87  | *64.38*  |
| M  | 1/24/2024  |   | 96.838  | 7.01  | -0.15  | 123.69  | 100.7  | 8.69  | *8.578*  | 22.9  | *22.8*  | 0.110  | 0.02  | 13.31  | *15.92*  | 0.01  | 53.62  | *64.00*  |
| F  | 2/21/2024  |   | 101.611  | 7.06  | -0.11  | 121.59  | 97.3  | 8.81  | *9.092*  | 20.0  | *19.9*  | 0.106  | 0.05  | 14.28  | *16.49*  | 0.08  | 58.30  | *66.36*  |
| A  | 3/15/2024  |   | 100.909  | 7.09  | 0.72  | 125.50  | 102.4  | 8.77  | *8.960*  | 22.8  | *22.9*  | 0.063  | 0.04  | 16.11  | *15.59*  | 0.21  | 64.62  | *62.73*  |
| K | 4/9/2024 |  | 101.367 | 7.00 | 0.06 | 125.41 | 100.3 | 8.58 | *8.578* | 22.9 | *23.0* | 0.067 | 0.03 | 16.26 | *15.62* | -0.03 | 64.83 | *62.84* |
| L | 5/7/2024 |  | 101.378 | 7.03 | 0.42 | 123.19 | 99.7 | 8.41 | *8.418* | 23.8 | *23.7* | 0.008 | -0.01 | 10.01 | *15.44* | 0.01 | 40.31 | *62.18* |
| M | 5/29/2024 |  | 100.912 | 7.01 | 1.12 | 122.93 | 99.5 | 8.27 | *8.341* | 24.5 | *24.7* | 0.045 | -0.03 | 14.66 | *15.17* | -0.02 | 58.94 | *61.17* |
| F | 6/27/2024 |  | 101.363 | 7.09 | 0.08 | 124.3 | 100.2 | 8.41 | *8.48* | 24.0 | *24.1* | 0.012 | -0.03 | 15.08 | *15.35* | -0.13 | 72.60 | *61.85* |
| K | 7/24/2024 |  | 100.941 | 7.04 | 1.32 | 118.78 | 99.6 | 8.25 | *8.263* | 24.9 | *25.1* | 0.089 | 0.17 | 11.51 | *15.32* | 0.57 | 46.54 | *61.74* |
| E | 8/22/2024 |  | 101.506 | 7.13 | 0.25 | 124.75 | 100.3 | 8.39 | *8.341* | 24.4 | *24.3* | 0.054 | -0.02 | 14.41 | *15.32* | -0.02 | 58.18 | *61.74* |
| M | 9/19/2024 |  | 101.471 | 7.07 | -0.11 | 124.36 | 100.4 | 8.39 | *8.341* | 24.3 | *24.3* | -0.014 | 0.01 | 15.37 | *15.29* | 0.03 | 61.64 | *61.63* |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |
| **BBMRRB**  | **Parameter**  | **Specific Conductivity (mS/cm)**  | **pH**  | **Turbidity (FNU)**  | **Dissolved Oxygen  (%, mg/L)**  | **Temperature (°C)**  | **Depth (m)**  | **Chlorophyll-a (RFU)**  | **Chlorophyll-a (ug/L)**  |
| **Sonde**  | **Date/Standard**  | **100**  | **7**  | **0**  | **124**  | **100%**  | **DO CCV**  | ***Std***  | **Sonde**  | ***Therm***  | **0**  | **0**  | **Chl CCV**  | ***Std***  | **0**  | **Chl CCV**  | ***Std***  |
| A  | 1/19/2024  | 100.142  | 6.97  | -0.11  | 120.51  | 99.9  | 8.62  | *8.661*  | 22.6  | *22.5*  | 0.033  | 0.01  | 13.23  | *15.92*  | -0.11  | 53.35  | *64.00*  |
| E  | 2/13/2024  | 100.678  | 6.98  | -0.14  | 124.30  | 98.8  | 8.44  | *8.578*  | 23.1  | *23.3*  | 0.022  | 0.02  | 15.07  | *15.56*  | -0.02  | 60.14  | *62.62*  |
| D  | 3/7/2024  | 101.047  | 7.08  | -0.03  | 124.44  | 100.1  | 8.56  | *8.578*  | 22.9  | *23.0*  | -0.004  | -0.06  | 17.80  | *15.83*  | -0.21  | 66.71  | *63.27*  |
| B | 4/4/2024 | 99.02 | 6.92 | 0.00 | 124.09 | 99.4 | 8.33 | *8.418* | 24.1 | *24.1* | -0.002 | 0.01 | 14.62 | *15.50* | 0.12 | 58.48 | *62.40* |
| A | 5/2/2024 | 101.180 | 6.98 | -0.08 | 123.71 | 102.7 | 8.60 | *8.418* | 24.1 | *24.2* | 0.03 | -0.02 | 20.90 | *15.38* | 0.06 | 80.51 | *61.96* |
| F | 5/24/2024 | 101.263 | 7.06 | 0.06 | 124.22 | 99.5 | 8.40 | *8.418* | 23.8 | *23.7* | 0.009 | 0.02 | 15.69 | *15.35* | 0.06 | 63.06 | *61.85* |
| G | 6/21/2024 | 101.091 | 7.08 | 0.02 | 126.49 | 100.4 | 8.44 | *8.418* | 24.0 | *24.0* | 0.066 | -0.03 | 14.91 | *15.44* | 0.06 | 57.57 | *62.18* |
| C | 7/11/2024 | 101.152 | 7.04 | 0.99 | 121.54 | 99.4 | 8.32 | *8.418* | 23.9 | *24.2* | 0.043 | N/A | N/A | *N/A* | N/A | N/A | *N/A* |
| A | 8/8/2024 | 101.132 | 7.01 | -0.19 | 123.82 | 99.4 | 8.26 | *8.341* | 24.5 | *24.9* | -0.016 | 0.03 | 14.90 | *15.17* | -0.14 | 59.88 | *61.17* |
| H | 9/6/2024 | 95.123 | 7.06 | 0.19 | 123.89 | 100.4 | 8.43 | *8.341* | 24.5 | *24.5* | -0.038 | 0.02 | 15.51 | *15.11* | 0.01 | 62.44 | *60.78* |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |
|  |
|  |
| **BBCWA4**  | **Parameter**  | **Specific Conductivity (mS/cm)**  | **pH**  | **Turbidity (FNU)**  | **Dissolved Oxygen  (%, mg/L)**  | **Temperature (°C)**  | **Depth (m)**  | **Chlorophyll-a (RFU)**  | **Chlorophyll-a (ug/L)**  |
| **Sonde**  | **Date/Standard**  | **100**  | **7**  | **0**  | **124**  | **100%**  | **DO CCV**  | ***Std***  | **Sonde**  | ***Therm***  | **0**  | **0**  | **Chl CCV**  | ***Std***  | **0**  | **Chl CCV**  | ***Std***  |
| H  | 1/19/2024  | 100.371  | 7.00  | -0.01  | 124.66  | 99.8  | 8.59  | *8.661*  | 22.6  | *22.7*  | 0.005  | -0.01  | 12.81  | *15.89*  | 0.00  | 51.37  | *63.75*  |
| J  | 2/13/2024  | 100.576  | 7.02  | -0.17  | 124.32  | 100.0  | 8.5  | *8.498*  | 23.3  | *23.3*  | 0.032  | -0.08  | 15.11  | *15.53*  | -0.26  | 60.46  | *62.51*  |
| K  | 3/7/2024  | 101.144  | 7.03  | 0.00  | 124.22  | 99.3  | 8.50  | *8.578*  | 23.0  | *23.0*  | -0.006  | -0.04  | 18.95  | *15.80*  | -0.03  | 69.51  | *53.50*  |
| L | 4/4/2024 | 101.300 | 6.92 | 0.12 | 124.08 | 100.0 | 8.38 | *8.418* | 24.1 | *24.1* | -0.002 | -0.01 | 17.87 | *15.50* | 0.01 | 71.67 | *62.40* |
| M | 5/2/2024 | 101.036 | 7.08 | -0.13 | 124.02 | 98.8 | 8.27 | *8.418* | 24.2 | *24.2* | 0.027 | 0.00 | 20.53 | *15.35* | 0.03 | 78.66 | *61.85* |
| J | 5/24/2024 | 100.807 | 7.02 | 0.04 | 123.46 | 99.0 | 8.35 | *8.418* | 23.8 | *23.7* | 0.011 | 0.09 | 14.56 | *15.26* | 0.22 | 58.02 | *61.52* |
| E | 6/21/2024 | 101.500 | 7.04 | 0.23 | 124.52 | 101.6 | 8.56 | *8.418* | 24.0 | *23.9* | 0.065 | 0.05 | 18.46 | *15.47* | 0.03 | 75.10 | *62.29* |
| L | 7/11/2024 | 100.747 | 7.01 | 0.34 | 123.04 | 98.30 | 8.24 | *8.418* | 24.2 | *23.9* | 0.039 | 0.03 | 15.34 | *15.11* | 0.02 | 61.75 | *60.91* |
| J | 8/8/2024 | 101.334 | 7.06 | -1.32 | 122.94 | 102.3 | 8.47 | *8.341* | 24.7 | *24.9* | 0.01 | 0.06 | 15.38 | *15.11* | 0.25 | 61.78 | *60.91* |
| N | 9/6/2024 | 101.944 | 6.97 | 0.08 | 124.08 | 110.5 | 9.13 | *8.341* | 24.5 | *24.5* | 0.024 | 0.03 | 15.01 | *15.14* | 0.00 | 60.41 | *61.04* |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |
|  |

**14) Other remarks/notes –**

Data are missing due to equipment or associated specific probes not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform. If additional information on missing data is needed, contact the Aquatic Preserve office. Copies of the calibration/deployment logs can be obtained through the Water Quality Specialist.

All data files are QAQC’d using the 2012 CDMO Excel macro. This macro automatically flags DO values less than 3 mg/L as Passed Initial QAQC Checks with the comment DO Hypoxia (<3 mg/L) (0, CDA). The DO flag is based on Federal standards. During the primary QAQC session, any reading that falls outside a sensor’s range (see section II. Physical Structure descriptors) is rejected (-3). Additionally, negative turbidity readings between (0) and (-2) have been automatically flagged as suspect data (1, CAF) since the value is within the accuracy of the probe per the CDMO operations manual. All negative chlorophyll readings are labeled as (-4) since they fall outside the chlorophyll sensor range and are manually rejected during QAQC (-3, SNV). All the first data points for each new deployment are labeled (0, CND).

The parameters that are QC’d are depth (m), temperature (°C), salinity (ppt), specific conductance (mS), turbidity (FNU), pH (SU), DO (mg/L and %) and chlorophyll (ug/L). All other parameters have not been QC’d.

All data collected after the 30-days in each deployment is noted in the F\_Record column of the dataset with {CSM}.

All turbidity readings over 124 FNU are rejected (-3, STS) since they are out of the calibration range that BBAP applies (0-124 FNU). All chlorophyll readings over the upper calibration value are rejected (-3, SCS) since they are out of the calibration range (varies by deployment). Additionally, for BBAP, anomalous turbidity and chlorophyll spikes are labeled suspect (1, STS/SCS) if BOTH of the following conditions apply:

1. The previous and next data points are BOTH less than the data point by the calculated value: STANDARD DEVIATION\*3 (SD calculated after rejected data removed)
2. Data point is not within a well-defined peak (e.g., there is no incremental increase before data point or decrease after data point based on the neighboring readings)

For BBLR03, turbidity values > 12 FNU from the previous and next data point were rejected (-3, STS), and values > 2 FNU from the previous and next data point were marked (1, STS). Chlorophyll values > STANDARD DEVIATION \*3 were rejected (-3, SCS), and values > 5.03 ug/L from the previous and next data point were marked (1, SCS).

For BBBB14, turbidity values > 7 FNU from the previous and next data point were rejected (-3, STS), and values > 6 FNU from the previous and next data point were marked (1, STS). Chlorophyll values > STANDARD DEVIATION\*3 were rejected (-3,SCS), ,and values > 2.78 ug/L from the previous and next data point were marked (1, SCS).

For BBJT71, turbidity values > 5 FNU from the previous and next data point were rejected (-3, STS), and values > 4 FNU from the previous and next data point were marked (1, STS). Chlorophyll values > STANDARD DEVIATION\*3 were rejected (-3,SCS), and values > 2.14 ug/L from the previous and next data point were marked (1, SCS).

For BBMRDW, turbidity values > 5 FNU from the previous and next data point were rejected (-3, STS), and values > 3 FNU from the previous and next data point were marked (1, STS). Chlorophyll values > STANDARD DEVIATION\*3 were rejected (-3, SCS), and values > 1.02 ug/L from the previous and next data point were marked (1, SCS).

For BBMRSC, turbidity values > 12 FNU from the previous and next data points were rejected (-3, STS), and values > 6 FNU from the previous and next data point were marked (1, STS). Chlorophyll values > STANDARD DEVIATION\*3 were rejected (-3, SCS), and values > 3.51 ug/L from the previous and next data point were marked (1, SCS).

For BBMRRB, turbidity values > 8 FNU from the previous and next data points were rejected (-3, STS), and values > 4 FNU from the previous and next data point were marked (1, STS). Chlorophyll values > STANDARD DEVIATION\*3 were rejected (-3, SCS) > 2.94 ug/L from the previous and next data point were marked (1, SCS).

For BBCWA4, turbidity values > 10 FNU from the previous and next data point were rejected (-3, STS), and values > 5 FNU from the previous and next data point were marked (1, STS). Chlorophyll values > STANDARD DEVIATION\*3 were rejected (-3,SCS), and values > 3.41 ug/L from the previous and next data point were marked (1, SCS).

These readings may either be caused by optical interference by animals or fouling, or an unknown local disturbance. Readings within a well-defined peak were maintained and given a (0) flag.

January 1-31, 2024

BBLR03

Note: 1/1 0:00 – 1/15 9:30, pH values significantly out of range, marked <-5>.

Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>), but marked (CDA). Example: 1/1 4:30 – 11:00.

Missing: 1/15 8:45-9:15, out for sonde exchange.

BBBB14

 Missing: 1/3 11:15 – 12:00; 1/30 9:30 – 9:45. Out for sonde exchange.

 Suspect: 1/10 8:00, 21:30; 1/11 7:15. Turbidity spikes out of range.

BBJT71

 Note: Data missing for 1/1, begins 1/2 but sonde did not record data and is marked <-2>.

 Suspect: 1/7 10:15. Chlorophyll spike out of range.

 Suspect: 1/21 8:30. Turbidity spike out of range.

 Missing: 1/30 10:15 – 10:30. Out for sonde exchange.

BBMRDW

Missing: 1/5 9:15 – 9:30. Out for sonde exchange.

Note: Sonde stopped recording data on 1/9 18:30 through the rest of the deployment. Cause unknown. New deployment began on 1/24 9:15.

BBMRSC

Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 1/1 0:00 – 1/4 7:30.

Missing 1/5 10:15 – 10:45; 1/24 9:15. Out for sonde exchange.

Note: Depth and salinity were recorded as “not a number” (NAN) and pH readings were significantly out of range from 1/16 7:45 through rest of the deployment. All parameters were marked (<-5>).

 BBCWA4

 Missing: 1/19 10:15 – 10:30. Out for sonde exchange

 Reject: 1/12 16:45. Turbidity spike out of range.

 Suspect: 1/14 2:45. Turbidity spike, cause unknown.

 BBMRRB

 Missing: 1/2 12:00 out for sonde exchange

Note: 1/13 - 1/14: Cold front, significant temperature drop

Reject: 1/15 14:00 – 17:15, 1/17 12:30- 20:15, 1/18 00:30 – 9:30 Turbidity spikes, possibly

due to blocked optics, fish found in sensor guard at end of deployment

Missing 1/30 10:15- 10:45: out for sonde exchange

February 1-28, 2024

BBLR03

Missing: 2/8 9:15 – 9:30 and 2/29 9:00 – 9:15, out for sonde exchange.

Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 2/20 23:30 – 2/21 6:00.

Suspect: 2/24 16:45, 17:15 turbidity spike, cause unknown.

Suspect 2/3 22:45, chlorophyll spike, cause unknown.

BBBB14

Suspect: 2/8 7:15, 9:15. Chlorophyll spike, cause unknown.

Note: Catastrophic sonde failure for deployment 2/26 – 3/21. No data available in that date range.

BBJT71

Reject: 2/9 12:45; 2/23 9:15; 2/26 15:30. Turbidity spike out of range.

Suspect: 2/13 12:00. Turbidity spike, cause unknown.

Note: DO hypoxia recorded 2/20 7:15. Data passed initial QAQC (<0>) but was marked (CDA).

Missing: 2/26 10:15. Out for sonde exchange.

BBMRDW

 Suspect: 2/4 21:30; 2/22 22:45. Turbidity spikes, cause unknown.

 Note: Sonde stopped recording data on 2/5 9:15 through the rest of the deployment. Cause unknown. New deployment began 2/21 9:45.

BBMRSC

Suspect: 2/2 19:30. Turbidity spike, cause unknown.

Reject: 2/4 21:30. Turbidity spikes out of range.

Note: Sonde stopped recording data on 2/5 9:15 through rest of the deployment. Cause unknown. New deployment began 2/21 10:45.

Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 2/24 2:15 – 4:30.

 BBCWA4

Missing: 2/13 11:15 – 12:00. Out for sonde exchange.

 Reject: 2/7 7:30. Chlorophyll spike out of range.

 Reject: 2/9 18:15; 2/14 1:15. Turbidity spikes out of range.

 Suspect: 2/28: Chlorophyll spike, cause unknown.

March 1-31, 2024

BBLR03

Missing: 3/28 8:15 – 8:30, out for sonde exchange.

Reject: 3/13 8:00. Turbidity spike out of range.

Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 3/24 00:45 – 6:30.

BBBB14

Note: Catastrophic sonde failure for deployment 2/26 – 3/21. No data available in that date range.

Note: Periodic extended readings of negative chlorophyll values. Data marked <-4>.

BBJT71

 Reject: 3/9 3:00; 3/20 18:30. Turbidity spikes out of range.

 Reject: 3/9 3:00; 3/11 11:45; 3/12 14:15. Chlorophyll spikes out of range.

 Missing: 3/21 8:45 – 9:00. Out for sonde exchange.

BBMRDW

 Missing: 3/15 8:30 – 9:00. Out for sonde exchange.

 Suspect: 3/23 13:45. Turbidity spike, cause unknown.

BBMRSC

Missing: 3/15 9:45 – 10:00. Out for sonde exchange.

Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 3/2 3:30 – 10:45.

BBMRRB

Suspect: Throughout the month, drops in salinity and Specific conductivity, and drops in

salinity and Specific conductivity with turbidity spike, possible blocked optic.

Turbidity and chlorophyll spikes occur throughout deployment, fish in guard or mount common at this site.

BBCWA4

 Reject: 3/2 12:45; 3/29 11:00; 3/31 16:30. Chlorophyll spikes out of range.

 Reject: 3/16 19:15; 3/18 13:45, 17:15; 3/19 6:15. Turbidity spikes out of range.

 Suspect: 3/20 7:30; 3/30 17:45. Turbidity spikes, cause unknown.

 Missing: 3/7 10:00 – 10:30. Out for sonde exchange.

April 1-30, 2024

 BBLR03

 Missing: 4/17 7:45-8:00. Out for sonde exchange.

 Suspect: 4/1 19:15-23:00. Chrlophyll spike, cause unknown.

 Reject: 4/12 6:00-7:15, 8:00-10:00; 4/13 3:30. Chlorophyll spikes out of range.

 BBBB14

 Missing: 4/15 9:45-10:00. Out for sonde exchange.

 Suspect: 4/25 17:45-19:00, 20:00-20:15. Turbidity spikes, cause unknown.

 BBJT71

 Missing: 4/15 9:00. Out for sonde exchange.

 BBMRDW

 Missing: 4/9 8:15-8:30. Out for sonde exchange.

 Suspect: 4/28 10:15. Turbidity spike, cause unknown.

 Note: pH sensor reading out of range and marked (-5) from 4/9 8:30 – 4/23 13:00.

 BBMRSC

 Missing: 4/9 9:30-9:45. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 4/20 10:00-13:00.

 Reject: 4/22 3:30; 4/28 1:00-1:15, 1:45-2:15, 5:45-6:00. Chlorophyll spikes out of range.

 BBMRRB

 Reject: 4/1 00:45-1:30, 2:00, 2:30-3:00, 3:30-4:45, 5:45-6:00, 19:15-19:45, 20:45, 21:30, 23:30- 23:45; 4/2 00:00-1:15, 1:45-3:30, 4:15-4:30, 5:30, 19:15, 20:15, 20:45-21:00, 21:30, 22:00, 23:30-23:45; 4/3 00:00-00:30, 1:15-3:00, 3:30-3:45, 4:30, 5:30-5:45, 21:00- 21:15, 22:00-22:15, 23:00, 23:30-23:45; 4/4 2:00-2:15, 2:45, 3:15, 3:45-4:00, 4:30, 5:00, 5:45, 19:30-19:45, 20:45; 4/5 5:00, 18:45; 4/6 2:00, 19:00, 20:15; 4/9 20:15; 4/10 19:00; 4/11 19:45; 4/12 11:00; 4/13 1:15-1:30, 18:00, 20:45; 4/14 7:00; 4/16 17:15, 18:00; 4/17 13:00, 23:15; 4/18 3:00, 14:15; 4/19 10:45; 4/20 11:45; 4/21 15:45, 19:30; 4/22 7:30, 12:00, 12:30, 21:30, 22:30; 4/23 7:15, 11:00, 12:30, 19:45, 22:00; 4/24 13:00-13:15, 21:00, 21:45; 4/25 1:00, 13:15; 4/27 1:30, 5:30, 8:30, 16:45, 17:15, 17:45; 4/28 2:00; 4/29 17:00, 17:30; 4/30 13:30, 15:00. Turbidity spikes out of range.

 Missing: 4/4 9:30-10:00. Out for sonde exchange.

 BBCWA4

 Missing: 4/4 8:45-9:00. Out for sonde exchange.

 Reject: 4/1 17:00; 4/4 6:30; 4/9 20:45, 22:00. Turbidity spikes out of range.

 Reject: 4/5 11:00; 4/7 00:45; 4/9 13:00, 13:30; 4/10 15:30-15:45, 16:30, 17:15; 4/11 18:00- 18:15; 4/20 13:15-14:00; 4/21 00:45 11:30-12:00, 12:30-12:45, 13:30, 23:45; 4/22 00:15-2:00; 4/26 13:15, 13:45; 4/27 2:45-4:00, 5:00. Chlorophyll spikes out of range.

May 1-31, 2024

 BBLR03

 Missing: 5/16 11:45-12:00. Out for sonde exchange.

 Suspect: 5/1 15:15. Turbidity spike, cause unknown

 Suspect: 5/27 16:15. Chlorophyll spike, cause unknown.

 Reject: 5/3 14:00; 5/15 16:00. Turbidity spike out of range.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 5/2 23:15 – 5/3 2:15.

 BBBB14

 Missing: 5/9 7:15-7:30. Out for sonde exchange.

 Note: No data recorded from 5/18 14:15 – 6/3 00:15 due to catastrophic sonde failure. Sonde inexplicably began recording data again on 6/3 00:30 until it was exchanged on 6/3.

 BBJT71

 Missing: 5/9 8:15-8:45. Out for sonde exchange.

 BBMRDW

 Missing: 5/7 8:30-9:00; 5/29 7:30-8:00. Out for sonde exchange.

 Reject: 5/3 13:30; 5/5 20:45. Turbidity spikes out of range.

 Suspect: 5/3 13:45; 5/6 15:15. Turbidity spike, cause unknown.

 Reject: 5/14 1:45, 11:30. Chlorophyll spikes out of range.

 BBMRSC

 Missing: 5/7 9:45-10:00; 5/29 9:00. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 5/2 23:30 – 5/3 8:00.

 Reject: 5/31 10:00. Turbidity spike out of range.

 Reject: 5/12 17:15; 5/14 18:00, 18:30, 19:15-19:30, 20:45; 5/21 5:30; 5/24 15:15. Chlorophyll spikes out of range.

 Suspect: 5/14 20:15; 5/24 15:00, 15:30; 5/25 14:30-15:30, 22:30. Chlorophyll spikes, cause unknown.

 BBMRRB

 Missing: 5/2 9:00-9:30; 5/24 8:15-11:15. Out for sonde exchange.

 Reject: 5/1 7:45, 11:00, 17:30, 18:00; 5/2 19:45, 20:45; 5/4 5:00, 20:45, 22:30, 23:45; 5/5 3:30-3:45, 4:30-5:00, 10:15, 20:30, 23:30; 5/6 00:45, 5:00, 13:00, 21:30, 22:00, 23:15; 5/7 14:45, 20:15, 20:45; 5/8 00:00-00:15, 1:30, 3:30, 4:00; 5/9 00:00-00:15, 20:00, 23:30; 5/10 1:45; 5/11 00:30, 3:30, 20:00, 20:30, 21:15; 5/12 00:00; 5/13 00:00, 1:00-1:15, 4:15, 19:30; 5/14 20:00; 5/16 2:00, 2:30, 20:45; 5/17 2:00, 4:00-4:15, 5:15 20:45; 5/18 00:00, 21:30; 5/19 00:45, 1:45, 11:15; 5/20 5:00-5:15. Turbidity spikes out of range.

 BBCWA4

 Missing: 5/2 8:00-8:30; 5/24 10:30-11:15. Out for sonde exchange.

 Reject: 5/8 20:00; 5/10 17:30; 5/18 18:15. Turbidity spikes out of range.

June 1-30, 2024

 BBLR03

 Missing: 6/25 8:45-9:00. Out for sonde exchange.

 Suspect: 6/11 2:30; 6/17 17:30-18:15; 6/18 17:45-20:15; 6/19 18:45-20:45; 6/20 7:00-8:00, 19:15-22:00 . Chlorophyll spike, cause unknown.

Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 6/13 3:00 – 6/17 17:15.

 BBBB14

 Missing: 6/5 8:15-8:30. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but

 Was marked (CDA). Example: 6/11 20:00 – 6/12 2:45.

 Suspect: 6/17 5:45 – 6/19 13:30. Chlorophyll spike out of range.

 BBJT71

 Missing: 6/5 8:45-9:00. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 6/11 14:45-23:00.

 Note: No data recorded from 6/28 18:45 through remainder of the month, cause unknown.

 Reject: 6/22 10:15. Turbidity spike out of range.

 Suspect: 6/17 20:00 – 6/19 13:00. Chlorophyll spike, cause unknown.

 BBMRDW

 Missing: 6/27 7:45-8:00. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 6/27 8:00-12:30.

 BBMRSC

 Missing: 6/27 9:00-9:15. Out for sonde exchange.

 Reject: 6/1 18:00; 6/4 23:00; 6/23 2:45, 6:00, 18:30, 19:00-19:45, 20:15; 6/27 3:15; 6/29 8:45. Chlorophyll spikes out of range.

 Suspect: 6/4 18:15; 6/5 13:45, 15:30, 17:30, 18:30, 20:15; 6/6 8:15; 6/23 20:00; 6/29 5:45. Chlorophyll spikes, cause unknown.

 BBMRRB

 Missing:

 Reject: 6/13 13:45. Turbidity spike out of range

 BBCWA4

 Missing: 6/21 8:30-8:45. Out for sonde exchange.

 Reject: 6/13 13:45; 6/25 7:15; 6/27 1:45; 6/28 23:45; 6/29 3:45. Turbidity spikes out of range.

July 1-31, 2024

 BBLR03

 Missing: 7/18 8:15-8:30. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 7/1 6:15-15:30.

 Reject: 7/18 13:30; 7/26 17:30. Turbidity spikes out of range.

 BBBB14

 Missing: 7/3 9:00-9:15. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 7/3 3:00-8:00.

 BBJT71

 Note: Data missing from 6/28 18:45 through 7/3 10:15. Cause unknown. Data recording continues as normal from thereon.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 7/6 22:45 – 7/7 4:45.

 BBMRDW

 Missing: 7/24 8:00-9:00. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 7/2 23:00 – 7/3 9:45

 Reject: 7/4 12:00. Turbidity spike out of range

 BBMRSC

 Missing: 7/24 10:30-11:00. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 7/2 8:45-18:00.

 Note: Data missing on 7/6 3:45-4:15. Cause unknown. Data resumes normal logging after 4:15 timestamp.

 Reject: 7/3 14:45. Turbidity spike out of range.

 Suspect: 7/3 13:15; 7/8 16:15. Turbidity spikes.

 Reject: 7/6 8:30; 7/8 5:00; 7/10 9:30-9:45, 12:45; 7/12 20:30, 21:30-22:00, 23:15-23:30; 7/13 8:30-9:00; 7/14 21:30, 22:00-23:15, 23:45 – 7/15 1:15; 7/15 6:15, 6:45, 8:15, 9:00, 9:45-10:15, 12:00-12:45, 13:15-13:30, 14:15, 19:30-19:45, 20:15-20:30; 7/16 3:15, 3:45-4:00, 5:30, 6:15-12:00, 12:30, 13:00-14:45, 15:15-15:30; 7/17 11:00, 12:30, 14:15 -14:30, 16:15; 7/17 22:00; 7/18 00:30, 1:30, 4:30, 5:00, 9:45; 7/19 00:15, 1:00, 2:00, 3:00, 23:00-23:15, 23:45; 7/20 00:00-4:00, 4:45-5:45; 7/22 00:45, 1:15, 1:45-3:15, 4:00-7:15, 7:45-8:45, 10:00, 12:30, 13:30-14:15, 14:45-16:15; 7/23 14:00. Chlorophyll spikes out of range.

 Suspect: 7/12 7:15-8:30. Chlorophyll spike.

 BBMRRB

 Note: Data missing from 7/1 00:00 through 7/11 11:45 due to sonde failure. The sonde has since been sent to YSI for repairs, but has not yet been returned. Data may yet be available when returned.

 BBCWA4

 Missing: 7/11 10:30-10:45. Out for sonde exchange.

 Reject: 7/11 14:15; 7/12 00:45, 1:30, 2:00-2:15, 5:45; 7/18 1:30, 20:00; 7/19 6:30; 7/20 17:30; 7/21 22:15; 7/25 19:30; 7/26 6:00; 7/29 17:45; 7/30 5:45-6:00, 15:45, 17:30; 7/31 5:45-6:15, 9:30. Turbidity spikes out of range.

 Reject: 7/1 22:15; 7/12 2:30, 20:15. Chlorophyll spikes out of range.

August 1-31, 2024

 BBLR03

 Missing: 8/15 8:00. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 8/2 00:30-5:15

 Reject: 8/2 17:45; 8/5 12:00, 13:30, 15:30; 8/12 10:00; 8/28 8:15. Chlorophyll spikes out of range.

 BBBB14

 Missing: 8/1 9:45-10:00; 8/27 9:00-9:15. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 8/1 4:15-8:30.

 Reject: 8/31 14:30, 21:00, 22:30, 23:00-23:15, 23:45. Turbidity spikes out of range.

 Note: 8/10 12:30; 8/11 15:00; 8/14 15:15; 8/24 2:15, 3:15; 8/27 2:00. Chlorophyll values outside of low sensor range and marked (<-4>).

 BBJT71

 Missing: 8/1 9:00; 8/27 9:45-10:00. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 8/10 3:45-8:15.

 Note: 8/27 10:15, 11:00-11:30, 12:00-14:45, 15:15-16:15, 16:45-18:00, 18:30, 19:00, 19:30- 22:15, 23:30-23:45; 8/28 00:15-1:15, 2:00-3:15, 4:00-4:30, 5:00-8:15, 9:00-9:30, 10:00 -11:15, 11:45-12:30, 13:00-13:15, 13:45-14:15, 14:45, 15:45, 17:30-17:45, 18:45 – 8/29 00:30, 1:00-3:15, 3:45-4:15, 5:00-17:15, 18:00, 19:15 – 8/30 6:15, 6:45-11:30, 12:00-14:15, 14:45-16:00, 16:30-17:15, 20:45-23:30; 8/31 00:15-2:00, 2:30-4:15, 4:45-5:15, 5:45-7:45, 9:30-10:30, 11:00-11:30, 12:00-15:00, 15:30-16:00, 20:30. Chlorophyll values outside of low sensor range and marked (<-4>).

 Suspect: 8/18 2:30. Turbidity spike.

 BBMRDW

 Missing: 8/22 9:00-9:45. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 8/10 6:45-10:30.

 Reject: 8/9 10:00. Turbidity spike out of range.

 BBMRSC

 Missing: 8/22 7:00-7:30. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 8/1 9:30-17:15.

 Reject: 8/1 11:15, 13:30, 14:00, 15:30-16:00, 18:30-18:45, 19:15-19:30, 20:15, 20:45, 22:00, 23:15; 8/6 19:00; 8/9 16:15, 17:45, 19:00; 8/10 2:45, 8:30, 13:45-14:15; 8/11 20:15; 8/12 9:00, 11:00, 13:15-14:00; 8/14 11:45, 12:45; 8/15 9:00; 8/17 6:15-6:30, 12:00 8/18 14:45; 8/19 20:00; 8/21 14:30. Chlorophyll spikes out of range.

 BBMRRB

 Missing: 8/8 8:30-9:00. Out for sonde exchange.

 Reject: 8/8 16:30; 8/19 5:45; 8/20 18:15; 8/23 15:00; 8/28 1:30. Turbidity spikes out of range.

 Suspect: 8/23 12:15. Turbidity spike.

 Suspect: 8/1 1:15; 8/4 3:45. Chlorophyll spikes.

 BBCWA4

 Missing: 8/8 9:45-10:00

 Suspect: 8/1 6:00, 7:15; 8/2 12:00; 8/4 18:00. Turbidity spikes.

 Reject: 8/2 6:45, 11:15-11:45, 21:15; 8/3 6:00; 8/4 6:00, 8:15; 8/5 8:30; 8/7 9:45; 8/8 00:30, 18:30; 8/11 19:00; 8/18 2:30, 6:30; 8/28 17:45. Turbidity spikes out of range.

 Suspect: 8/1 7:00; 8/6 8:00. Chlorophyll spikes.

 Reject: 8/5 11:30; 8/10 18:15; 8/27 17:30. Chlorophyll spikes out of range.

September 1-30, 2024

 BBLR03

 Missing: 9/10 7:45-8:00. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 9/1 22:30 – 9/2 7:30.

 Reject: 9/30 13:45-14:45. Turbidity spikes out of range.

 BBBB14

 Missing: 9/24 9:30-9:45. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 9/10 20:00 – 9/11 3:30.

Reject: 9/1 00:00, 00:30, 1:00, 2:15-2:30, 3:30, 4:15-4:30, 5:15, 6:45, 16:15, 19:30-19:45; 9/2 19:15, 21:00; 9/3 00:45, 1:15, 3:15; 9/5 23:15; 9/17 6:45; 9/22 20:00, 22:45-23:00, 23:30-23:45; 9/23 00:00, 00:30, 1:00-1:30; 9/27 21:30. Turbidity spikes out of range.

 Suspect: 9/2 11:15, 21:45; 9/5 13:45; 9/22 23:15. Turbidity spikes out of range.

 Reject: 9/2 23:30; 9/19 20:15; 9/22 4:45; 9/26 15:00. Chlorophyll spikes out of range.

 BBJT71

 Missing: 9/24 8:45-9:00. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 9/24 5:15-7:45.

 Reject: 9/23 13:15. Turbidity spike out of range.

 Note: Periodic extended periods of chlorophyll values outside of low sensor range and marked (<-4>). Example: 9/2 13:00-14:45.

 BBMRDW

 Missing: 9/19 9:15-9:30. Out for sonde exchange.

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 9/1 2:45-9:30.

 Suspect: 9/9 20:45. Chlorophyll spike out of range.

 BBMRSC

 Note: Periodic extended periods of DO hypoxia. Data passed initial QAQC (<0>) but was marked (CDA). Example: 9/4 1:45-10:00.

 Reject: 9/9 18:45. Turbidity spike out of range.

 Reject: 9/12 8:45; 9/21 5:30; 9/22 19:15, 20:00-20:15; 9/23 3:45, 8:45; 9/25 21:30; 9/29 8:45-9:00, 12:00; 9/30 11:45-12:00, 15:00. Chlorophyll spikes out of range.

 BBMRRB

 Missing: 9/6 9:00-9:15. Out for sonde exchange.

 BBCWA4

 Missing: 9/6 9:45. Out for sonde exchange.

 Reject: 9/5 6:45; 9/6 17:30; 9/7 17:45. Turbidity spikes out of range.

 Suspect: 9/11 7:45. Turbidity spike.

 Reject: 9/18 6:00; 9/24 12:00, 22:15-22:45; 9/25 15:45-16:00; 9/26 15:45. Chlorophyll spikes out of range.

**15) Acknowledgement:**

The data included with this document were collected by the staff of the Florida Department of Environmental Protection at Biscayne Bay Aquatic Preserves. Any products derived from these data should clearly acknowledge this source (please use the attached logos). This recognition is important for ensuring that this long-term monitoring program continues to receive the necessary political and financial support.