Biscayne Bay Aquatic Preserves (BBAP) Water Quality Metadata

January – July 2025

Latest Update: July 25, 2025

Note: This is a provisional metadata document; it has not been authenticated as of its download date.  Contents of this document are subject to change throughout the QAQC process and it should not be considered a final record of data documentation until that process is complete.  Contact the Aquatic Preserve office general email at Biscayne.Bay@FloridaDEP.gov with any additional questions.

I.  Data Set and Research Descriptors

1)  Principal investigator and contact persons –

Tristan Clute – Water Quality Technician

(786)-798-4515, Tristan.Clute@FloridaDEP.gov

1277 NE 79th St, Miami, FL 33138

- Datasonde specialist, telemetry coordinator, water quality data analyst, report preparer, Quality Assurance/ Quality Control (QAQC) officer, Field and lab logistics manager, equipment calibration & deployment coordinator, data handling and upload coordinator, cross-network lab coordinator

Griffin Alexander - Biscayne Bay Aquatic Preserves Manager

(786) 798-4511, Griffin.N.Alexander@FloridaDEP.gov

1277 NE 79th St., Miami FL, 33138

- Project supervisor, field assistance

TBD - Water Quality Specialist

- data handling, field assistance, datasonde calibration & deployment, report preparer

2)  Entry verification –

Deployment data are uploaded from the YSI datasonde to a DEP issued Lab laptop.  Files are exported from EcoWatch in a comma-delimited format (.CDF), EcoWatch Lite in a comma separated file (.CSV),  KOR EXO v2.3.10.0 , or KOR v1.1.8.0 Software in a comma separated file (.CSV) and uploaded to the National Estuarine Research Reserve System (NERRS) Centralized Data Management Office (CDMO) Non-System Wide Monitoring Program (SWMP) Data Upload Service where data undergo automated primary QAQC. All pre- and post-deployment data are removed from the file prior to upload.

During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Office of Resilience and Coastal Protection (RCP) Data Coordinator and/or the Aquatic Preserve office for secondary QAQC where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheet for flagged data and summary statistics, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, remove any overlapping deployment data, append files, and export the resulting data file for upload to the Aquatic Preserve (AP) database.

Upload after secondary QAQC results in incorporation into the AP database as provisional plus data, and finally tertiary QAQC by the RCP’s Data Coordinator and assimilation into the AP database as authenticated data.  Where deployment overlap occurs between files, the data produced by the newly calibrated datasonde is accepted as being the most accurate. For more information on QAQC flags and codes, see Sections 11 and 12.

Data was handled by Claire Burgett from January 2019 - January 2020, by Aliza Karim from February 2020 - May 2022, Ellyn Willse from May 2022 – July 2023. Christopher Mayer from July 2023 to December 2024, and CURRENTLY by Tristan Clute from December 2024 – present.

3)  Research objectives –

This three-step research project was initiated in response to a 2,000-3,000-acre seagrass loss event in Northern Biscayne Bay, specifically in and around the Julie Tuttle Basin (Basin). Stations were chosen in an array to understand the relationship between the Little River flow into the Basin and across the large, central shoal containing seagrass habitat. Stations with historical data (not necessarily datasonde data) sampled by other agencies were chosen for this project, when appropriate. Not all parameters are taken at all sites across the project, however some stations included sampling for all three research steps. Four sonde stations are set up down the Miami River and across the Rickenbacker Basin to monitor flow from the Miami River across the basin and towards the Bill Sadowski Critical Wildlife Area.

- Currently, seven deployed datasonde stations monitor abiotic parameters at 15-minute intervals.

- In addition to the datasondes, this project involves taking monthly water quality discrete surface samples at 23 sites with an additional 2 discrete 1 meter-from-bottom samples taken at the furthest up-river sampling sites, LR03, and MRSC. The water quality discrete sampling tests for 41 analytes including nutrients, pharmaceuticals, herbicides, fungicides, insecticides, and chlorophyll-a and include sampling sites at all datasonde stations.

- The 34 benthic assessment sites include seagrass and macroalgae tissue sampling for elemental analyses and stable isotopes, sediment sampling & depth, Braun-Blanquet & cover abundance surveys, and some abiotic factors. All 34 sites are sampled semiannually with several datasonde stations having benthic sites within close proximity.

4)  Research methods –

YSI datasondes are deployed monthly at seven permanent locations and one temporary site including BBLR03, BBBB14, BBJT71, BBMRSC, BBMRDW, BBMRRB, BBCWA4 and the temporary site of BBVS01. Data is collected in 15-minute intervals and all sondes are calibrated before deployment and a pre-deployment calibration verification (ICV) is done as a QAQC check. Approximately every month, the deployed sondes are removed from the water and returned to the lab for data retrieval, cleaning, verifying post-deployment calibration (CCV) and conducting any necessary maintenance or repairs. Freshly calibrated sondes are swapped at the time of retrieval.

All time is reported as Eastern Standard Time.

YSI 6600 and 6920 datasondes have been phased out but were operating in the Bay until April 15th, 2020. Currently, YSI EXO2 (10m depth) sondes are swapped between each station. The sonde is placed into a horizontally deployed PVC pipe with holes to allow for water flow. The unit is mounted to a concrete slab or dock piling and submerged.

BBAP was granted EPA funding to expand the water quality program into the Miami River and Rickenbacker basin in 2022. The seagrass in this basin showed resilience to the die-off event record in 2019 but there is more stakeholder/researcher concerns due to the outflow of water through the Government Cut inlet and onto the coral reef tracks. With this EPA funding, we established two sites in the Miami River in March 2021 and two sites in the Rickenbacker basin in January 2022.

In the Miami River, downstream of a tributary that is known to release nutrient/metal pollutants, an EXO2 site was established on March 31st, 2021 (site: BBMRDW). The mount is attached to a dock and held at depth via an anchor where the sonde is retrieved by lifting the PVC mount out of the water.  At the salinity control structure of the Miami River, which is about 9 miles upriver, an EXO2 site was established on March 26th, 2021 (site: BBMRSC). The mount is attached to a concrete seawall and the sonde is retrieved by lifting the PVC mount out of the water.

Near the mouth of the Miami River, EXO2 site BBMRRB was established January 13th, 2022. The sonde is placed in a vertically deployed mount that is constructed from PVC (with flow through holes) attached to an FWC owned piling with custom brackets. EXO2 Site BBCWA4 is located on the edge of the Bill Sadowski Critical Wildlife Area in the same basin as BBMRRB and was established January 18th, 2022.  The sonde is placed in a vertically deployed mount that is constructed from PVC (similar to the BBMRRB site)  The locations for sites BBMRSC, BBMRDW, BBMRBB and BBCWA4 follow the path of water down the Miami River and out into the Bay depending on tidal cycles.

Specific conductance, pH, turbidity, depth, chlorophyll and dissolved oxygen are calibrated on the EXO2s while temperature is calibrated in factory settings. Calibrations are conducted in the BBAP lab by BBAP staff within 24hours of deployment.

Specific conductance is 1-point calibrated with the 50,000 uS/cm standard (from Ricca). The ICV is done with the 1,000 uS/cm standard while the CCV is done with the 100,000 uS/cm standard. pH is 2-point calibrated with the pH7 standard and pH10 standard (from Fisher Scientific). The ICV and CCV are done with one of the 2 standards. Turbidity is 2-point calibrated with 0 FNU/NTU Milli-Q water and 124 FNU standard (from YSI). The ICV and CCV are done with both the 0 and the 124/126 FNU/NTU standard. Chlorophyll is 2-point calibrated with 0 ug/L Milli-Q water and a rhodamine WT dye standard (from Kingscote). The Tal-PE sensor on the EXO 2 is calibrated for chlorophyll-a (ug/L and RFU).  Chlorophyll calibrations using rhodamine WT dye allow for consistency in values across our sondes and deployments. During monthly water quality sampling, sonde readings are paired with water bottle grabs that are tested for chlorophyll using a fluorometric method, allowing for verification of sonde chlorophyll readings.

5)  Site location and character –

All sites are in relatively close proximity with tidal range around 2 feet and part of the same watershed. The watershed is an urban residential part of Miami-Dade County. There are fertilizer concerns and septic tank leaching considerations flowing into the watershed with repeated issues involving sewage overflows and other spills.

BBLR03 is in between the mouth of the Little River and the salinity control structure and is heavily tidally influenced. The salinity control structures hold back the Bay’s tidal wedge that would otherwise travel upriver possibly leading to damaged systems and aquifer. Once filled these structure open for scheduled releases coinciding with the outgoing tide where a rush of fresh water is pushed out from the bottom of the structures as the gates are slightly lifted upwards. Little River, as the pulses in freshwater occur, is suspected to be a major source of nutrients and other pollutants into the Basin and generally has poorer water quality than the bay sites. This site was part of Miami-Dade County’s long-term water quality monitoring BayRun program but monitoring at this location was suspended until we began sampling in 2019. Freshwater inflow into Little River and therefore the bay can radically change conditions and drop DO levels but is determined by the salinity control structure and varies depending on management actions, rainfall, and timed releases of Lake Okeechobee water from the salinity control structures.

- Location: 25.846841, -80.182861

- Salinity range: 0.0 – 35.0 ppt

- Depth: 0.01 – 1.48 m (it’s about 10 feet in center of river; sonde is up against the seawall along a private dock)

- Bottom habitat: bare muddy sediment, little to no vegitation

- Pollutants detected in 2023:

- Pollutants detected in 2024:

BBBB14 is in the Northern Biscayne Bay Basin in the northwest section of the shoal. This site was highly impacted by the seagrass die-off event in the northern portion of the shoal where full loss of the prior seagrass cover was observed. The seagrass loss event occurred as an edge-in effect suggesting locations nearest Little River could have been first impact of dieoff event . The site, BBBB14, is part of Miami-Dade County’s long-term water quality monitoring BayRun program and was likely one of the first sites to switch to a macroalgal dominated bed. Benthic sampling by BBAP at this site began in 2019 and involves cardinal direction quadrat throws. It’s a brackish site with freshwater inflow primarily coming from rainfall and secondarily from the nearby Little River but also has exchange with other basins in Biscayne Bay influenced by other rivers and canals.  Seagrass had not been observed in quadrats during our 2019 benthic surveys at this site, however extremely sparse Syringodium filiforme was seen in March 2020. Historical benthic data was not taken at this station, however sites nearby showed high seagrass cover prior to the die-off event. Now, into July of 2025 the Syringodium seagrass beds of Julia Tuttle basin have largely recovered and are showing signs of Thalassia intrusion with sparse baren or algal dominated regions nearest river mouths.

- Location: 25.83003, -80.15860

- Salinity range: 14.6 – 36.0 ppt

- Depth: 0.12 – 2.0 m

- Bottom habitat: Currently dominated by Syringodium filiforme. Sediment is largely made up of soft mud and Halimeda hash (remains of calcified algal discs) at the surface in a muddy matrix.

- Pollutants detected in 2023:

- Pollutants detected in 2024:

BBJT71 – Decommissioned March 27, 2025 -  is in the Northern Biscayne Bay Basin in the southeast section of the shoal, further from the Little River than BBBB14. This site was chosen to observe the area where some seagrass still survived or potentially had not yet been impacted by the die-off event at the initiation of the project. The southeast interior section of the shoal in the Basin was previously fully covered by dense refugia of seagrasses, with the dominant species being Syringodium filiforme. Historical data showed averages of 50-75% Syringodium filiforme cover before the die-off event, 1% Syringodium filiforme after die off event, but now are seeing quadrats with 40%< seagrass cover. Historical data at this site is primarily composed of benthic surveys, but there was a short term datasonde deployment by county researchers. This site was sampled originally by United States Geological Survey (USGS) during Fish and Invertebrate Assessment Network surveys in 2014. Miami-Dade County then resampled this location following the seagrass die-off event. Our program established a transect in 2018 at this location for quarterly benthic surveys.
After 5+ years of data collection the BB14 site and this one were shown to be redundant, so this site was removed in March of 2025.

- Location: 25.821730, -80.151250

- Salinity range: 24.2 – 35.9 ppt

- Depth: 0.42 – 2.24 m

- Bottom habitat: Currently dominated Syringodium and sparsely by other macroalgal species. Sediment is muddy sand with some Halimeda hash (remains of calcified algal discs).

- Pollutants detected in 2023:

- Pollutants detected in 2024:

BBMRDW is located 1.5-miles upriver from the mouth of the Miami River, downriver of the intersection with Wagner Creek (suspected to be a major source of nutrients and pollutants such as heavy metals). The Miami River outflows into the Rickenbacker Basin and then out the Government Cut inlet which eventually leads to the offshore coral reef tract. The site is tidally influenced with stratification of waters becoming more predominant in wet season. Water quality in the area is poor with nutrient sources coming from urban run-off and development consisting of regular low dissolved oxygen events. There is also significant amount of cargo ship and charter vessel traffic through the Miami River.

- Location: 25.7773, -80.2040

- Salinity range: 0.7 – 34.2 ppt

- Depth: 0.35 – 1.92 m

- Bottom habitat: bare muddy sediment with some hard substrate lining the seawall

- Pollutants detected in 2023

- Pollutants detected in 2024:

BBMRSC is located about 9.1-miles upriver from the mouth of the Miami River, directly bellow the outflow of the C-6 salinity control structure with potential nutrient and pollutant sources likely from urban run-off. The Miami River outflows into the Rickenbacker Basin and then out the Government Cut inlet which leads to the offshore coral reef tract. The site is tidally influenced with stratification often seen in wet season versus well mixed conditions in the dry season. Freshwater inflow from the salinity control structure varies depending on management actions, rainfall, and timed releases of Lake Okeechobee water but consistently reads in the hypoxic range especially as the structures are initially opened.

- Location: 25.8072, -80.2599

- Salinity range: 0.2 – 26.2 ppt

- Depth: 0.89-2.48 m

- Bottom habitat: bare muddy sediment

- Pollutants detected in 2023

  - Surface:

  - Bottom

- Pollutants detected in 2024:

  - Surface:

  - Bottom:

BBCWA4 is located on the Rickenbacker Basin shoal on the western edge of the Bill Sadowski Critical Wildlife Area. This site is tidally influenced and receives water from offshore via Government cut inlet and freshwater input from the Miami River. This site is adjacent to an area of dense seagrass habitat. This area is of concern as the invasive seagrass species Halophila Stipulacea appears to be well established in the more sheltered portions of this protected area and has been seen growing even epiphytically on the sonde mount itself.

- Location: 25.758020, -80.167690

- Salinity range: 27.6-37.6

- Depth: 3.07-1.93

- Bottom habitat: mixed seagrass and sponge habitat, near dense seagrass

- Pollutants detected in 2023

- Pollutants detected in 2024

BBMRRB is located in the Rickenbacker Basin at the confluence of the Miami River and the bay. This area receives heavy boat traffic. This site was selected to monitor water flowing out of the Miami River as it makes its way into the basin and offshore via Government cut inlet. Benthos is often silted but the paddle grass H. decipiens was found near station.

- Location: 25.769800, -80.180850

- Salinity range: 25.9-37.3 ppt.

- Depth: 2.65-1.40 m

- Bottom habitat: very sparse seagrass such as H. decipiens, fair amount of green & brown algas, spots of bare sediment with harder almost packed muddy sediment

- Pollutants detected in 2023

- Pollutants detected in 2024

BBAP Station Deployment Timeline:

| Station Code | Station Name | Location | Active Dates | Notes |
| --- | --- | --- | --- | --- |
| BBLR03 | Little River (LR03) | 25.846841, -80.182861 | 4/19/2019 - present | Model 6920 (Jan 1-Mar 17, 2020), EXO2 (Mar 17, 2020-present) |
| BBBB14 | Tuttle Basin NW (BB14) | 25.830030, -80.158600 | 1/18/2019 - present | Model 6600V2-4 (Jan 1-Apr 15, 2020), EXO2 (Apr 15, 2020-present) |
| BBJT71 | Tuttle Basin SE (JT71) | 25.821730, -80.151250 | 3/5/2019 – 3/27/2025 | Model EXO2 |
| BBMRDW | Miami River (MRDW) | 25.7773, -80.204 | 3/31/2021 – present | Model EXO2 |
| BBMRSC | Miami River (MRSC) | 25.8072, -80.2599 | 3/25/2021 - present | Model EXO2 |
| BBMRRB | Rickenbacker Basin (MRBB) | 25.76980, -80.18085 | 1/13/2022 - present | Model EXO2 |
| BBCWA4 | Rickenbacker Basin (CWA4) | 25.758020, -80.167690 | 1/18/2022 - present | Model EXO2 |

6)  Data collection period – EST

BBBB14

| Deployment date/time | Retrieval date/time | Notes |
| --- | --- | --- |

12/18/2024 11:30                     	1/23/2025 12:00

1/23/2025 12:15	2/20/2025 9:45

2/20/2025 10:00	3/20/2025 8:30

3/20/2025 8:45	4/15/2025 8:30

4/15/2025 8:45	5/9/2025 9:30

5/9/2025 9:45 	6/13/2025 8:45

6/13/2025 9:00 	7/9/2025 10:30

7/9/2025 10:45 	8/8/2025 8:45

8/8/2025 9:00 	9/9/2025 10:45

9/9/2025 11:00 	10/9/2025 8:30

10/9/2025 8:45 	11/13/2025 11:15

11/13/2025 11:30 	12/11/2025 11:15

12/11/2025 11:30 	1/6/2026 13:15

BBJT71

| Deployment date/time | Retrieval date/time | Notes |
| --- | --- | --- |

12/18/2024 10:15                     1/23/2025 11:45

1/23/2025 11:45		      2/20/2025 10:15

2/20/2025 10:15 		      3/20/2025 10:15

3/20/2025 10:15 		      3/26/2025 10:30	Decommissioned

BBLR03

| Deployment date/time | Retrieval date/time | Notes |
| --- | --- | --- |

12/5/2024 9:30	            1/3/2025 12:30

1/3/2025 12:45 	            2/5/2025 13:30

2/5/2025 13:45                        3/05/2025 12:15

3/05/2025 12:30	            4/03/2025 11:15

4/3/2025 11:45 	            5/6/2025 11:30

5/6/2025 11:45 	            6/5/2025 12:00

6/5/2025 12:15 	            7/1/2025 10:15

7/1/2025 10:30 	            7/29/2025 10:00

7/29/2025 10:15 	            9/2/2025 10:30

9/2/2025 10:45	            10/2/2025 10:00

10/2/2025 10:15	            11/4/2025 10:45

11/4/2025 11:00 	            11/25/2025 10:45

11/25/2025 11:00 	            12/23/2025 10:45

12/23/2025 11:00 	            1/8/2026 11:30

BBMRDW

| Deployment date/time | Retrieval date/time | Notes |
| --- | --- | --- |

12/11/2024 9:45	                   1/7/2025 10:15

1/7/2025 10:45 		     2/5/2025 10:45

2/5/2025 11:00		     3/05/2025 10:45

3/05/2025 10:45		     4/03/2025 9:30

4/03/2025 9:45		     5/06/2025 9:30

5/06/2025 9:45		     6/05/2025 9:15

6/05/2025 9:30		     7/01/2025 9:15

7/1/2025 8:30 		     7/29/2025 9:15

7/29/2025 9:45		     9/2/2025 9:30

9/2/2025 9:45  		     10/2/2025 9:00

10/2/2025 9:30 		     11/4/2025 10:00

11/4/2025 10:15 	     11/25/2025 10:00

11/25/2025 10:15 	     12/23/2025 10:00

12/23/2025 10:15 	     1/8/2026 10:30

BBMRSC

| Deployment date/time | Retrieval date/time | Notes |
| --- | --- | --- |

12/11/2024 11:00 	     1/7/2025 13:00

1/7/2025 13:00 		     2/5/2025 11:30

2/5/2025 12:15		     3/5/2025 11:30

3/5/2025 11:30 		     4/3/2025 10:30

4/3/2025 10:45 		     5/6/2025 10:15

5/6/2025 10:45 		     6/5/2025 11:00

6/5/2025 11:15 		     7/1/2025 8:30

7/1/2025 9:00 		     7/29/2025 8:00

7/29/2025 8:15		     9/2/2025 9:15

9/2/2025 9:30		     10/2/2025 11:15

10/2/2025 9:45 		     11/4/2025 9:00

11/4/2025 9:15 		     11/25/2025 9:30	Depth sensor failure – no data recorded

11/25/2025 9:30 		     12/23/2025 9:15

12/23/2025 9:30 		     1/8/2026 9:30

BBMRRB

| Deployment date/time | Retrieval date/time | Notes |
| --- | --- | --- |

12/20/2024 11:00 	           1/23/2025 10:30

1/23/2025 10:30	           2/20/2025 12:00

2/20/2025 12:15 	           3/20/2025  11:00

3/20/2025 11:30 	           4/15/2025 10:30

4/15/2025 10:45 	           5/9/2025 11:30

5/9/2025 11:45 	           6/13/2025 11:15

6/13/2025 11:30 	           7/9/2025 9:00

7/9/2025 9:15 	           8/8/2025 10:45

8/8/2025 11:00	           9/9/2025 8:30

9/9/2025 9:45	           10/9/2025 10:30

10/9/2025 10:45 	           11/13/2025 12:45	Batteries died no data collected after

11/13/2025 13:00 	           12/10/2025 16:00	C/T sensor cracked during deployment – most data rejected

12/10/2025 16:15 	           1/6/2026 10:45

BBCWA4

| Deployment date/time | Retrieval date/time | Notes |
| --- | --- | --- |

12/20/2024 12:00 	     1/23/2025 10:00

1/23/2025 10:00		     2/20/2025 12:45

2/20/2025 12:45		     3/20/2025 12:00

3/20/2025 12:30 		     4/15/2025 11:00

4/15/2025 11:15 		     5/9/2025 12:30

5/9/2025 12:45 		     6/13/2025 10:30

6/13/2025 10:45 		     7/9/2025 8:30

7/9/2025 8:45 		     8/8/2025 10:00

8/8/2025 10:00		     9/9/2025 8:45

9/9/2025 9:00 	 	     10/9/2025 10:00

10/9/2025 10:15 		     11/13/2025 13:45	C/T sensor error – most data rejected

11/13/2025 14:00 	     12/10/2025 15:45	Batteries died no data collected after 11/2

12/10/2025 16:00 	     1/6/2026 11:15

7)  Distribution –

The Principle Investigator (PI) retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the Florida Department of Environmental Protection and the Aquatic Preserve site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.

Aquatic Preserve water quality data and metadata can be obtained from the Manager at the individual Aquatic Preserve site (please see Principal Investigators and Contact Persons) and online at the Aquatic Preserve’s data portal home page www.floridaapdata.org.  Data are available in comma delimited format.

8)  Associated researchers and projects –

There are several agencies conducting water quality monitoring within Biscayne Bay. The following are the projects that we are aware of though there may be more.

(This Project) Florida Department of Environmental Protection (DEP) – Biscayne Bay Aquatic Preserves (BBAP)

- Currently, eight deployed datasonde stations monitor abiotic parameters at 15-minute intervals.

- In addition to the datasondes, this project involves taking monthly water quality grab samples at 23 sites. The 41 analytes are tested from water quality discrete grab sampling, include nutrients, pharmaceuticals, herbicides, fungicides, insecticides, and chlorophyll-a. All datasonde stations are also monthly water quality grab sampling stations. Data can be acquired on the Aquatic Preserves AP Data Portal website.

- The 34 benthic assessment sites include seagrass and macroalgae tissue sampling for elemental analyses and stable isotopes, sediment sampling & depth, Braun-Blanquet & percent cover abundance surveys, and some abiotic factors. Previously quarterly sampling of eleven benthic sites and all twenty-one sites semiannually occurred. Currently there are 34 sites which are monitored semi-annually.

  - Sediment sulfide analyses have also been done in conjunction with the SAV sampling at all sites in September 2020.

  - In April 2021, grain size analyses were conducted in the Rickenbacker basin and sedimentary iron sampling was done at all benthic sites in both basins.

- https://floridadep.gov/rcp/aquatic-preserve/locations/biscayne-bay-aquatic-preserves

- BBAP has conducted (from August 2020 to present) monthly bird surveys on 4 islands within Northern Biscayne Bay in response to a bird rookery collapse on Bird Key island. Data can be acquired from Griffin Alexander.

Florida Department of Environmental Protection (DEP) – Office of Resilience and Coastal Protection (ORCP)

- ORCP is handling the establishment and characterization of the Government Cut Inlet Contributing Area (ICA) that includes the upland watershed that flows into Biscayne Bay, Biscayne Bay itself, and the coastal reef area just offshore of the government cut inlet. The watershed will be characterized by partners at NOAA via increased water quality monitoring in collaboration with this BBAP project, Miami-Dade County’s BayRun project, and NOAA’s pollutant loading model development. Methodology development conversations are ongoing with a planned initiation in fall 2021. The program’s goals include addressing EPA’s 9 elements of a comprehensive Watershed Management Plan for the ICA.

- Further research will be initiated through a DEP Coral Reef Protection and Restoration Grant program that will be awarded in late 2020/early 2021

Florida International University (FIU)

- Datasonde monitoring via the Center for Aquatic Chemistry and Environment (CREST) program with live telemetry data available in sites located in the same basin and adjacent basins as the BBAP sondes available through Dr. Piero Gardinali.

- https://crestcache.fiu.edu/research/research-buoys/index.html

- Long-term water quality sampling within Biscayne Bay as part of the Southeast Environmental Research Center (SERC) Water Quality Monitoring Network.

- http://serc.fiu.edu/wqmnetwork/

- EPA funded project on hydrodynamic modeling of Biscayne Bay with preliminary results available through Dr. Henry Briceno.

- FIU contract via the seagrass lab through Dr. Jim Fourqurean to assess BBAP collected samples of sediments and seagrass and algae tissues.

- Various graduate student projects within and around Biscayne Bay.

Miami-Dade County Department of Environmental Resources Management (DERM)

- Long-term monitoring for environmental data and water quality parameters within Biscayne Bay, canals, and their tributaries (BayRun program).

- Data available via request and through .pdf reports.

- http://www.miamidade.gov/environment/surface-water-quality.asp

South Florida Water Management District (SFWMD)

- Long-term surface water quality monitoring at coastal and canal sites in central and south Biscayne Bay.

- Flow monitoring from canal and river outputs around Biscayne Bay.

- Groundwater, sediment, rainfall and weather monitoring data available as well.

- Data stored and accessible through DBHYDRO database.

- http://my.sfwmd.gov/dbhydroplsql/show_dbkey_info.main_menu

Miami Waterkeeper (MWK) and Florida Department of Health (DOH)

- DOH Healthy Beaches program supplemented by bay sampling from Miami Waterkeeper

- Sample monitoring for fecal indicator bacteria.

- http://www.floridahealth.gov/environmental-health/beach-water-quality/index.html

- Continuous surface parameter data available via deployed buoys in Northern Biscayne Bay.

- https://crestcache.fiu.edu/research/research-buoys/index.html

National Oceanic and Atmospheric Administration (NOAA)

- Long-term water quality sampling.

- Turbidity study done in Northern Biscayne Bay, including a station in the Tuttle Basin that included instrument deployments and bottle grabs.

- https://pdfs.semanticscholar.org/70c8/c68607f1bf40bc13c331fb2d8e09213f5658.pdf

University of Florida, Institute of Food and Agricultural Sciences (IFAS) Sea Grant Extension Program

- Biscayne Bay Water Watch Program.

- Community-based volunteer monitoring program to take over SFMWD and DERM sites that lost funding.

- Abiotic, nutrient, and chlorophyll a data collected.

- https://sfyl.ifas.ufl.edu/miami-dade/natural-resources/biscayne-bay-water-watch-/

University of Central Florida

- eDNA analysis of surface waters at 8- ecologically different sites with Biscayne Bay (4 of which will be in Northern Biscayne Bay). Project to be initiated in Aug. 2020 by Dr. Michelle Gaither’s Lab.

- https://sciences.ucf.edu/biology/gaitherlab/

- Rapid Evaluation of Estrogenic Endocrine Disrupting Compounds sampling event through surface sampling at 10 sites in Biscayne Bay. Project to be initiated Fall 2020 by Dr. John Fauth’s Lab.

University of Miami

- Substrate and temporal viability of enterococci bacteria by Dr. Helena Solo-Gabriele and students.

- Various graduate student projects within and around Biscayne Bay.

- https://www.rsmas.miami.edu/research/projects/index.html

Kent State University

- Imaging and analysis of fluorescent pigment data collected through satellite imagery by Dr. Joseph Ortiz.

- In preliminary stages but intended to be able to determine the spatial extent of seagrass versus macro-algal dominated beds as well as chlorophyll blooms.

II.  Physical Structure Descriptors

9)  Sensor specifications –

Currently, we have 13 YSI EXO2 Sondes that are interchangeably swapped monthly at each station. At BBLR03, prior to March 17, 2020, two YSI 6920s were swapped monthly. At BBBB14, prior to April 15, 2020, one YSI 6600 V2-4 was deployed monthly.

YSI EXO2 (10m depth) data sonde:

Parameter: Total algae

Units: micrograms/Liter

Sensor Type: optical sensor

Model#: YSI 599103-01

Range: 0 to 400 µg/L chl

Accuracy: Linearity: r2  ≥ 0.999 for Rhodamine WT across full range

Resolution: 0.01 RFU or 0.01 µg/L chl

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: 4-electrode nickel cell

Model#: YSI 599827

Range:0 to 200 mS/cm

Accuracy: 0 to 100: +/- 0.5% of reading or 0.001 mS/cm, w.i.g.; 100 to 200 +/- 1% of reading

Resolution: 0.0001 to 0.01 mS/cm

Parameter: Temperature

Units: Celsius (C)

Sensor Type: 4-electrode nickel cell

Model#: YSI 599827

Range: -5 to 35 ºC, 35 to 50 ºC

Accuracy: -5 to 35 ºC: +/- 0.01 ºC2, 35 to 50 ºC: +/- 0.05 ºC2

Resolution: 0.001 ºC

Parameter: dissolved oxygen, % air saturation

Units: percent air saturation (%)

Sensor Type: optical sensor

Model#: YSI 599100-01

Range: 0 to 500% air saturation

Accuracy: 0 to 200%: +/1 1% of reading or 1% saturation, w.i.g.; 200 to 500% +/- 5% of reading

Resolution: 0.1% air saturation

Parameter: dissolved oxygen, mg/L

Units: mg/L

Sensor Type: optical sensor

Model#: YSI 599100-01

Range: 0 to 50 mg/L

Accuracy: 0 to 20 mg/L: +/- 0.1 mg/L or 1% of reading, w.i.g.; 20 to 50 mg/L: 5% of reading

Resolution: 0.01 mg/L

Parameter: pH

Units: pH units

Sensor Type: unguarded

Model#: YSI 599702

Range: 0 to 14 units

Accuracy: +/- 0.1 pH units within +/- 10ºC of calibration temp; +/- 0.2 pH units for entire temp range

Resolution: 0.01 units

Parameter: Turbidity

Units: FNU (formazin nephelometric units)

Sensor Type: optical sensor

Model#: YSI 599101-01

Range: 0 to 4000 FNU

Accuracy: 0 to 999 FNU: 0.3 FNU or +/- 2% of reading, w.i.g.; 1000 to 4000 FNU: +/- 5% of reading

Resolution: 0 to 999 FNU=0.01 FNU; 1000 to 4000 FNU=0.1 FNU

Parameter: Depth

Units: meters

Sensor Type: integral, non-vented depth sensor

Range: 0 to 10 m

Accuracy: +/- 0.04 FS (+/- 0.004m or +/-0.013ft)

Resolution: 0.001 ft (0.001 m)

Sensor Disclaimers:

Depth Qualifier:

YSI data sondes can be equipped with either vented or non-vented depth/level sensors.  Readings for both vented and non-vented sensors are automatically compensated for water density change due to variations in temperature and salinity; but for all non-vented depth measurements, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.02 cm for every 1 millibar change in atmospheric pressure and is eliminated for vented sensors because they are vented to the atmosphere throughout the deployment time interval.

Standard calibration protocol calls for all non-vented depth sensors to read 0 meters at a (local) barometric pressure of 1013.25 mb (760 mm/hg).  To achieve this, each site calibrates their depth sensor with a depth offset number, which is calculated using the actual atmospheric pressure at the time of calibration and the equation provided in the Aquatic Preserve calibration sheet or digital calibration log.  This offset procedure standardizes each depth calibration. If accurate atmospheric pressure data are available, non-vented sensor depth measurements can be corrected. The Principal Investigator should be contacted in order to obtain information regarding atmospheric pressure data availability.

Salinity Units Qualifier:

The 6600 series sondes report salinity in parts per thousand (ppt) units, the EXO sondes report practical salinity units (psu). These units are essentially the same and for Aquatic Preserve purposes are understood to be equivalent, however psu is considered the more appropriate designation. Moving forward the Aquatic Preserve program will assign psu salinity units for all data regardless of sonde type.

Turbidity Qualifier:

The 6600 series sondes report turbidity in nephelometric turbidity units (NTU), the EXO sondes use formazin nephelometric units (FNU). These units are essentially the same but indicate a difference in sensor methodology, for Aquatic Preserve purposes they will be considered equivalent.  Moving forward, the Aquatic Preserve program will use FNU/NTU as the designated units for all turbidity data regardless of sonde type. If turbidity units and sensor methodology are of concern, please see the Sensor Specifications portion of the metadata.

Chlorophyll Fluorescence Disclaimer:

YSI chlorophyll sensors (6025 or 599102-01) are designed to serve as a proxy for chlorophyll concentrations in the field for monitoring applications and complement traditional lab extraction methods; therefore, there are accuracy limitations associated with the data that are detailed in the YSI manual including interference from other fluorescent species, differences in calibration method, and effects of cell structure, particle size, organism type, temperature, and light on sensor measurements.

10)  Coded variable definitions –

Sampling station:		Station code:

LR03			BBLR03

JT71			BBJT71

BB14			BBBB14

MRSC			BBMRSC

MRDW			BBMRDW

CWA4			BBCWA4

MRRB			BBMRRB

VS01			BBVS01

11)  QAQC flag definitions –

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is missing and above or below sensor range.  All remaining data are then flagged 0, passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3		Data Rejected due to QAQC

-2		Missing Data

-1		Optional SWMP Supported Parameter

0		Data Passed Initial QAQC Checks

1		Suspect Data

2		Open - reserved for later flag

3		Calculated data: non-vented depth/level sensor correction for changes in barometric pressure

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12)  QAQC code definitions –

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the deployment or YSI datasonde, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an * below) can be applied to the entire record in the F_Record column.

General Errors

GIC	No instrument deployed due to ice

GIM	Instrument malfunction

GIT	Instrument recording error; recovered telemetry data

GMC 	No instrument deployed due to maintenance/calibration

GNF	Deployment tube clogged / no flow

GOW	Out of water event

GPF	Power failure / low battery

GQR	Data rejected due to QA/QC checks

GSM	See metadata

Corrected Depth/Level Data Codes

GCC	Calculated with data that were corrected during QA/QC

GCM	Calculated value could not be determined due to missing data

GCR	Calculated value could not be determined due to rejected data

GCS	Calculated value suspect due to questionable data

GCU 	Calculated value could not be determined due to unavailable data

Sensor Errors

SBO	Blocked optic

SCF	Conductivity sensor failure

SCS		Chlorophyll spike

SDF	Depth port frozen

SDG	Suspect due to sensor diagnostics

SDO	DO suspect

SDP	DO membrane puncture

SIC		Incorrect calibration / contaminated standard

SNV	Negative value

SOW	Sensor out of water

SPC	Post calibration out of range

SQR	Data rejected due to QAQC checks

SSD	Sensor drift

SSM	Sensor malfunction

SSR		Sensor removed / not deployed

STF	Catastrophic temperature sensor failure

STS		Turbidity spike

SWM	Wiper malfunction / loss

Comments

CAB*	Algal bloom

CAF	Acceptable calibration/accuracy error of sensor

CAP	Depth sensor in water, affected by atmospheric pressure

CBF	Biofouling

CCU	Cause unknown

CDA*	DO hypoxia (<3 mg/L)

CDB*	Disturbed bottom

CDF	Data appear to fit conditions

CFK*	Fish kill

CIP	*	Surface ice present at sample station

CLT*	Low tide

CMC*	In field maintenance/cleaning

CMD*	Mud in probe guard

CND	New deployment begins

CRE*	Significant rain event

CSM*	See metadata

CTS	Turbidity spike

CVT*	Possible vandalism/tampering

CWD*	Data collected at wrong depth

CWE*	Significant weather event

13)  Post deployment information –

CCV values in red did not pass the established acceptance criteria for the Florida AP database. Data from the deployments that did not pass CCV were labeled as suspect. AC was ±5% of standard value for specific conductance, turbidity, DO (mg/L) & pHmV with ±2% for chlorophyll and BGA, ±3% for DO%, ±0.2 for pH and temperature (°C). CCV values in dark red that are bold, fell beyond 3x the acceptance criteria so the data from the deployment was rejected. Values in italics are standard values. Standard values for DO (mg/L), Chlorophyll (ug/L), Chlorophyll (RFU) and BGA (RFU) are determined by temperature.

| BBAP | Son-de | Depl Date mm/dd | Spc Cond. mS/cm | pH |  | Turbidity FNU |  | Dissolved Oxygen %, mg/L |  |  | Temperature °C |  | Depth m |  | Chlorophyll-a ug/L |  |  |  | Chlorophyll-a RFU |  |  | BGA RFU |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Station | A-M | 2025 | 100k or 1k | 7 / 10 |  | DI / 124 |  | % / mg/L / Standard |  |  | Sonde / Therm. |  | See Offset Chart |  | DI / Mix / Standard |  |  |  | DI / Mix / Standard |  |  | DI / Mix / Standard |  |  |
| BBLR03 | M | 1/03 | 100.84 | 6.90 |  | 0.48 | 125.17 | 101.5 | 8.93 | 8.829 | 21.6 | 21.5 | 0.076 | -0.03 |  | 65.58 | 64.86 |  |  |  |  | - | - | - |
| BBLR03 | B | 2/06 | 101.088 | 7.07 | 10.13 | 0.19 | 124.15 | 102.00 | 8.88 | 8.661 | 23.2 | 23.3 | 0.108 | 0.10 |  | 64.40 | 63.88 |  |  |  |  | - | - | - |
| BBLR03 | D | 3/05 | 101.201 | 6.98 | 10.04 | 0.15 | 128.18 | 99.7 | 8.62 | 8.661 | 22.6 | 22.5 | -0.015 | -0.19 |  | 74.33 | 73.50 |  |  |  |  | -0.61 | - | 43.00 |
| BBLR03 | L | 4/03 | 101.399 | 6.96 | 10.01 | 0.10 | 128.37 | 101.40 | 8.68 | 8.578 | 23.10 | 23.059 | 0.107 | -0.01 |  | 60.53 | 62.40 |  |  |  |  | -0.27 | 37.87 | 43.00 |
| BBLR03 | I | 5/06 | 102.402 | 7.01 | 9.95 | 0.23 | 128.12 | 101.9 | 8.62 | 8.498 | 23.60 | 23.64 | 0.070 | 0.01 |  | - | - | -0.18 |  | - | - | -2.26 | - | - |
| BBLR03 | J | 6/05 | 98.001 | 7.05 | 10.05 | -0.06 | 123.46 | 99.7 | 8.38 | 8.418 | 24.048 | 24.1 | 0.067 | -0.21 |  | 70.03 | 61.52 | 0.03 |  | 16.72 | 15.26 | -9.35 | 58.28 | 41.00 |
| BBLR03 | E | 7/01 | 100.776 | 7.11 | 10.09 | 0.00 | 125.00 | 107.60 | 8.36 | 7.759 | 24.472 | 24.5 | 0.078 | -0.05 |  | 65.18 | 62.40 | -0.13 |  | 15.76 | 15.50 | -0.18 | 40.59 | 41.00 |
| BBLR03 | L | 7/29 | 100.936 | 6.99 | 10.00 | 0.39 | 124.95 | 101.10 | 8.43 | 8.341 | 24.4 | 24.5 | -0.088 | 0.09 |  | 58.94 | 62.07 | 0.03 |  | 14.58 | 15.41 | -0.28 | 31.49 | 41.00 |
| BBLR03 | N | 9/02 | 100.706 | 7.03 | 9.94 | 0.30 | 122.60 | 101.6 | 8.46 | 8.341 | 24.564 | 24.6 | 0.038 | 0.14 |  | 55.63 | 62.62 | 0.04 |  | 13.80 | 15.56 | 0.37 | 35.48 | 41.00 |
| BBLR03 | I | 10/1 | 101.344 | 7.03 | 10.02 | 0.61 | 124.44 | 100.60 | 8.53 | 8.498 | 23.500 | 23.6 | 0.116 | 0.07 |  | 60.73 | 63.75 | 0.04 |  | 15.09 | 15.86 | 0.29 | 38.34 | 43.00 |
| BBLR03 | A | 11/4 | 1.002 | 7.11 | 10.04 | -0.10 | 124.92 | 101.20 | 8.65 | 8.578 | 23.172 | 23.2 | 0.092 | -0.09 |  | 63.28 | 65.13 | -0.01 |  | 16.13 | 15.74 | -0.55 | 41.53 | 41.00 |
| BBLR03 | N | 11/24 | 98.655 | 6.98 | 10.07 | 0.52 | 122.32 | 104.1 | 8.76 | 8.578 | 23.033 | 23.1 | 0.137 | 0.02 |  | 59.21 | 63.28 | 0.01 |  | 15.02 | 15.74 | -0.02 | 41.62 | 43.00 |
| BBLR03 | F | 12/23 | 100.133 | 7.04 | 10.03 | 0.06 | 124.27 | 100.5 | 8.70 | 8.636 | 22.471 | 22.5 | 0.095 | -0.10 |  | 64.79 | 63.39 | 0.00 |  | 15.96 | 15.77 | -0.20 | 43.34 | 43.000 |
| BBBB14 | G | 1/23 | 104.102 | 7.03 |  | 0.41 | 123.93 | 101.6 | 9.70 | 9.566 | 17.4 | 17.4 | 0.091 | 0.08 |  | 70.24 | 69.72 |  |  |  |  | - | - | - |
| BBBB14 | L | 2/21 | 101.420 | 7.05 | 10.06 | 0.04 | 124.65 | 102.9 | 9.08 | 8.829 | 21.2 | 21.2 | 0.111 | -0.06 |  | 65.19 | 64.50 |  |  |  |  | -0.32 | 38.05 | 45.00 |
| BBBB14 | C | 3/20 | 101.681 | 7.00 | 10.07 | 0.35 | 126.85 | 101.80 | 9.00 | 8.829 | 21.408 | 21.4 | 0.017 | 0.11 |  | 65.06 | 63.28 |  |  |  |  | -0.19 | 21.65 | 43.00 |
| BBBB14 | K | 4/16 | 97.363 | 7.16 | 10.09 | 0.28 | 127.47 | 100.80 | 8.67 | 8.578 | 22.8 | 22.8 | 0.068 | 0.30 |  | 62.50 | 64.25 |  |  |  |  | -0.26 | 46.150 | 43.00 |
| BBBB14 | C | 5/09 | 0.994 | 7.19 | 10.18 | 0.11 | 124.19 | 101.2 | 8.08 | 8.041 | 26.783 | 26.9 | 0.072 | 0.16 |  | - | - | 0.03 |  | - | - | 1.26 | - | - |
| BBBB14 | I | 6/13 | 100.746 | 6.99 | 10.01 | -0.04 | 123.99 | 102.9 | 8.62 | 8.418 | 24.148 | 24.1 | 0.069 | -0.34 |  | 63.77 | 61.63 | 0.00 |  | 15.69 | 15.29 | -0.50 | 41.85 | 41.00 |
| BBBB14 | J | 7/09 | 100.139 | 7.08 | 10.10 | 0.09 | 124.62 | 101.30 | 8.53 | 8.498 | 23.853 | 23.9 | 0.045 | 0.09 |  | 58.89 | 62.95 | 0.06 |  | 14.57 | 15.62 | -2.73 | 36.90 | 41.00 |
| BBBB14 | I | 8/07 | 101.350 | 6.99 | 10.01 | 0.11 | 124.45 | 101.00 | 8.47 | 8.418 | 24.111 | 24.1 | 0.047 | -0.13 |  | 55.26 | 63.50 | -0.02 |  | 14.01 | 15.80 | -0.87 | 36.17 | 41.00 |
| BBBB14 | H | 9/08 | 100.475 | 7.00 | 10.15 | 0.19 | 125.70 | 100.00 | 8.41 | 8.418 | 23.998 | 24.0 | -0.018 | 0.04 |  | 57.83 | 63.75 | -0.01 |  | 14.87 | 15.86 | -0.01 | 39.86 | 41.00 |
| BBBB14 | M | 10/9 | 100.864 | 6.94 | 9.88 | 0.11 | 123.92 | 101.5 | 8.93 | 8.829 | 21.625 | 21.6 | 0.103 | 0.11 |  | 65.80 | 65.25 | 0.03 |  | 16.26 | 16.22 | -0.16 | 40.07 | 41.00 |
| BBBB14 | H | 11/13 | 101.442 | 6.94 | 9.96 | -0.11 | 123.39 | 101.3 | 8.81 | 8.743 | 22.167 | 22.1 | 0.037 | -0.09 |  | 60.21 | 64.13 | -0.01 |  | 15.67 | 15.95 | -0.59 | 39.90 | 41.00 |
| BBBB14 | G | 12/9 | 100.599 | 6.99 | 9.98 | 0.14 | 122.80 | 100.1 | 8.59 | 8.578 | 22.953 | 23.0 | 0.025 | -0.13 |  | 66.79 | 64.25 | -0.08 |  | 16.41 | 15.98 | -0.42 | 44.18 | 43.00 |
| BBJT71 | K | 1/23 | 101.87 | 7.01 |  | 0.21 | 121.64 | 100.6 | 9.59 | 9.566 | 17.4 | 17.4 | 0.110 | -0.15 |  | 71.46 | 69.72 |  |  |  |  | - | - | - |
| BBJT71 | J | 2/20 | 101.065 | 7.05 | 10.01 | 0.39 | 124.64 | 102.10 | 9.00 | 8.829 | 21.2 | 21.2 | 0.106 | 0.24 |  | 66.97 | 64.75 |  |  |  |  | -4.92 | 63.55 | 43.00 |
| BBJT71 | N | 3/20 | 101.770 | 7.08 | 10.16 | 0.236 | 126.58 | 101.4 | 8.96 | 8.829 | 21.4 | 21.4 | 0.027 | 0.10 |  | 64.33 | 63.88 |  |  |  |  | 0.01 | 22.95 | 43.00 |
| BBJT71 | G | 3/27 | 101.226 | 7.09 | 10.08 | 0.39 | 125.62 | 100.10 | 8.52 | 8.498 | 23.2 | 23.2 | 0.082 | -0.08 |  | 61.22 | 64.13 |  |  |  |  | -0.01 | 41.52 | 43.00 |
| BBMRRB | E | 1/24 | 102.352 | 6.98 |  | 0.29 | 124.46 | 101.40 | 8.68 | 8.566 | 17.4 | 17.3 | 0.045 | -0.06 |  | 71.30 | 69.72 |  |  |  |  | - | - | - |
| BBMRRB | M | 2/20 | 100.548 | 7.04 | 10.06 | 0.22 | 124.24 | 102.20 | 9.02 | 8.829 | 21.2 | 21.2 | 0.110 | -0.18 |  | 64.84 | 65.38 |  |  |  |  | -0.10 | 38.26 | 43.00 |
| BBMRRB | B | 3/20 | 101.648 | 6.99 | 10.07 | 0.45 | 126.94 | 101.7 | 8.98 | 8.829 | 21.4 | 21.4 | 0.027 | 0.34 |  | 65.16 | 64.88 |  |  |  |  | -0.14 | - | - |
| BBMRRB | A | 4/16 | 99.912 | 6.97 | 10.03 | 0.23 | 126.27 | 100.90 | 8.69 | 8.661 | 22.8 | 22.7 | 0.091 | 0.09 |  | 63.12 | 63.17 |  |  |  |  | -0.26 | 48.44 | 43.00 |
| BBMRRB | N | 5/09 | 1.049 | 7.20 | 10.20 | 66.46 | 125.50 | 100.3 | 8.00 | 8.041 | 26.914 | 26.7 | 0.067 | 1.07 |  | - | - | 0.29 |  | - | - | 0.05 | - | - |
| BBMRRB | D | 6/11 | 100.785 | 7.02 | 10.07 | -0.83 | 125.74 | 102.1 | 8.58 | 8.418 | 24.070 | 24.1 | 0.061 | 0.09 |  | 65.46 | 62.40 | -0.02 |  | 16.20 | 15.50 | 0.96 | 60.87 | 41.00 |
| BBMRRB | B | 7/09 | 101.261 | 7.07 | 10.07 | -0.30 | 124.93 | 100.90 | 8.498 | 8.49 | 23.875 | 23.9 | 0.034 | 0.01 |  | 68.47 | 62.29 | -0.04 |  | 17.09 | 15.50 | -0.29 | 39.51 | 41.00 |
| BBMRRB | F | 8/07 | 101.493 | 7.05 | 10.10 | 0.26 | 125.34 | 100.50 | 8.418 | 8.44 | 24.026 | 24.1 | 0.036 | -0.14 |  | 55.72 | 63.06 | -0.02 |  | 14.05 | 15.68 | -1.20 | 36.20 | 41.00 |
| BBMRRB | L | 9/08 | 101.453 | 6.99 | 10.05 | -0.47 | 118.57 | 100.20 | 8.43 | 8.418 | 23.974 | 24.0 | -0.014 | -0.69 |  | 56.65 | 62.29 | 0.01 |  | 13.71 | 15.47 | -0.15 | 38.02 | 43.00 |
| BBMRRB | N | 10/7 | 100.924 | 7.00 | 10.06 | 0.06 | 124.54 | 102.00 | 8.97 | 8.829 | 21.601 | 21.6 | 0.093 | -0.06 |  | 66.02 | 64.75 | -0.02 |  | 16.65 | 16.10 | 0.09 | 40.41 | 41.00 |
| BBMRRB | F | 11/13 | 100.644 | 6.94 | 10.00 | 0.14 | 12.3.02 | 101.60 | 8.85 | 8.743 | 22.094 | 22.1 | 0.085 | -0.13 |  | 61.23 | 65.00 | -0.02 |  | 15.96 | 16.16 | -0.20 | 39.74 | 41.00 |
| BBMRRB | C | 12/9 | 100.862 | 7.03 | 10.04 | 0.25 | 122.91 | 100.00 | 8.57 | 8.578 | 22.993 | 23.0 | 0.023 | -0.21 |  | 67.47 | 64.63 | -0.01 |  | 16.77 | 16.07 | 0.62 | 43.71 | 43.00 |
| BBCWA4 | D | 1/24 | 97.001 | 6.99 |  | 0.16 | 124.82 | 100.90 | 9.63 | 9.566 | 17.4 | 17.4 | 0.035 | -0.02 |  | 69.70 | 69.72 |  |  |  |  | - | - | - |
| BBCWA4 | A | 2/20 | 99.965 | 7.06 | 10.04 | -0.01 | 123.73 | 102.10 | 9.01 | 8.829 | 21.1 | 21.2 | 0.118 | -0.31 |  | 65.04 | 65.63 |  |  |  |  | 1.12 | 10.99 | 43.00 |
| BBCWA4 | H | 3/20 | 101.450 | 7.19 | 10.18 | 0.31 | 126.10 | 102.30 | 9.02 | 8.829 | 21.5 | 21.4 | 0.209 | 0.29 |  | 64.02 | 65.00 |  |  |  |  | -0.15 | - | 43.00 |
| BBCWA4 | I | 4/16 | 98.843 | 7.05 | 10.08 | 0.20 | 126.70 | 103.9 | 8.93 | 8.578 | 22.8 | 22.8 | 0.081 | 0.10 |  | 63.30 | 63.50 |  |  |  |  | -1.68 | 47.44 | 43.00 |
| BBCWA4 | E | 5/09 | 1.023 | 7.10 | 10.15 | 0.36 | 124.78 | 101.00 | 8.11 | 8.041 | 26.506 | 26.5 | 0.078 | 0.34 |  | - | - | 0.11 |  | - | - | -0.15 | - | - |
| BBCWA4 | K | 6/12 | 100.995 | 7.04 | 10.06 | 0.04 | 122.68 | 102.20 | 8.58 | 8.418 | 24.124 | 24.1 | 0.067 | 0.30 |  | 62.16 | 62.60 | 0.09 |  | 15.36 | 15.56 | 5.76 | 63.98 | 41.00 |
| BBCWA4 | N | 7/09 | 101.345 | 7.03 | 9.93 | -0.05 | 124.90 | 100.40 | 8.46 | 8.498 | 23.877 | 23.9 | 0.019 | -0.06 |  | 62.61 | 62.07 | 0.00 |  | 15.36 | 15.41 | -0.74 | 38.14 | 41.00 |
| BBCWA4 | C | 8/07 | 101.086 | 7.03 | 10.10 | 0.01 | 129.08 | 101.40 | 8.51 | 8.418 | 24.089 | 24.1 | 0.042 | -0.06 |  | 55.93 | 63.28 | 0.01 |  | 14.42 | 15.74 | -0.47 | 37.41 | 41.00 |
| BBCWA4 | A | 9/08 | 101.217 | 6.98 | 9.99 | -1.08 | 130.27 | 101.40 | 8.53 | 8.418 | 23.964 | 24.0 | -0.012 | 0.22 |  | 64.60 | 62.84 | 0.07 |  | 15.85 | 15.62 | 0.25 | 40.44 | 41.00 |
| BBCWA4 | K | 10/9 | 56.390 | 6.98 | 9.78 | -0.05 | 123.87 | 85.10 | 5.93 | 8.829 | 34.834 | 21.6 | 0.092 | -0.07 |  | 65.35 | 64.63 | -0.02 |  | 16.17 | 16.07 | 0.19 | 39.63 | 41.00 |
| BBCWA4 | E | 11/13 | 100.994 | 6.96 | 9.99 | 0.31 | 123.66 | 100.7 | 8.78 | 8.743 | 22.072 | 22.1 | - | 0.00 |  | 60.21 | 64.13 | 0.03 |  | 15.67 | 15.95 | 0.31 | 41.35 | 41.00 |
| BBCWA4 | L | 12/9 | 101.069 | 6.97 | 9.99 | 0.22 | 122.71 | 100.2 | 8.60 | 8.578 | 22.893 | 22.9 | 0.043 | 0.01 |  | 67.22 | 64.88 | 0.01 |  | 16.55 | 16.16 | -0.14 | 42.76 | 43.00 |
| BBMRDW | A | 1/07 | 101.310 | 7.00 |  | 0.29 | 124.46 | 101.40 | 8.68 | 8.566 | 17.4 | 17.3 | 0.068 | 0.46 |  | 62.21 | 64.75 |  |  |  |  | - | - | - |
| BBMRDW | F | 2/06 | 101.229 | 6.97 | 9.98 | -0.26 | 121.03 | 102.8 | 8.83 | 8.661 | 23.2 | 23.3 | 0.103 | 3.40 |  | 56.94 | 63.39 |  |  |  |  | - | - | - |
| BBMRDW | K | 3/05 | 95.274 | 6.94 | 10.03 | 0.01 | 125.87 | 100.80 | 8.71 | 8.661 | 22.6 | 22.6 | -0.022 | 0.03 |  | 73.84 | 73.50 |  |  |  |  | 0.02 | - | - |
| BBMRDW | M | 4/03 | 101.197 | 7.00 | 9.96 | -0.01 | 126.82 | 102.6 | 8.77 | 8.578 | 23.1 | 23.1 | 0.098 | 0.02 |  | 63.76 | 63.06 |  |  |  |  | -0.18 | 41.60 | 43.00 |
| BBMRDW | K | 5/06 | 101.251 | 6.94 | 10.01 | 0.30 | 121.52 | 100.20 | 8.49 | 8.498 | 23.626 | 23.6 | 0.059 | 0.16 |  | - | - | 0.05 |  | - | - | -1.94 | - | - |
| BBMRDW | M | 6/05 | 100.051 | 7.04 | 10.13 | -0.01 | 125.31 | 101.50 | 8.49 | 8.341 | 24.256 | 24.3 | 0.062 | -0.26 |  | 67.14 | 62.07 | 0.00 |  | 15.93 | 15.41 | -7.44 | 61.14 | 41.00 |
| BBMRDW | F | 7/01 | 101.459 | 7.18 | 10.16 | 4.28 | 104.40 | 101.20 | 7.85 | 7.759 | 24.504 | 24.6 | 0.072 | -0.46 |  | 48.51 | 61.96 | -0.08 |  | 11.91 | 15.38 | 9.23 | 36.83 | 41.00 |
| BBMRDW | G | 7/29 | 101.059 | 7.15 | 10.11 | 0.20 | 132.22 | 105.60 | 8.80 | 8.341 | 24.492 | 24.6 | -0.008 | -0.08 |  | 62.49 | 61.52 | -0.01 |  | 15.47 | 15.26 | 0.12 | 33.03 | 41.00 |
| BBMRDW | J | 9/02 | 0.011 | 7.09 | 10.01 | 0.27 | 124.29 | 100.10 | 8.31 | 8.341 | 24.823 | 24.7 | 0.040 | 0.03 |  | 56.74 | 61.96 | 0.06 |  | 14.26 | 15.41 | -0.21 | 33.78 | 41.00 |
| BBMRDW | F | 10/2 | 101.586 | 7.02 | 10.14 | 0.59 | 124.40 | 104.80 | 8.90 | 8.498 | 23.500 | 23.4 | 0.106 | 0.07 |  | 59.77 | 64.25 | 0.02 |  | 14.87 | 15.71 | 0.09 | 37.64 | 43.00 |
| BBMRDW | L | 11/4 | 0.998 | 7.06 | 10.04 | 0.07 | 125.06 | 100.3 | 8.61 | 8.578 | 23.191 | 23.2 | 0.998 | 0.26 |  | 63.94 | 63.28 | 0.06 |  | 15.91 | 15.74 | 0.41 | 41.44 | 41.00 |
| BBMRDW | J | 11/24 | 99.604 | 7.03 | 10.04 | 0.28 | 124.42 | 104.20 | 8.77 | 8.578 | 23.121 | 23.1 | 0.121 | 0.31 |  | 64.04 | 64.00 | 0.11 |  | 15.68 | 15.92 | 0.29 | 43.80 | 43.00 |
| BBMRDW | E | 12/23 | 100.387 | 6.97 | 9.99 | 0.11 | 124.46 | 99.5 | 8.60 | 8.661 | 22.557 | 22.5 | 0.061 | 0.10 |  | 65.68 | 63.63 | -0.01 |  | 16.12 | 15.83 | 0.98 | 44.95 | 43.00 |
| BBMRSC | L | 1/07 | 101.870 | 7.01 |  | 0.27 | 124.38 | 99.8 | 8.75 | 8.829 | 21.6 | 21.7 | 0.067 | 0.09 |  | 63.03 | 64.25 |  |  |  |  | - | - | - |
| BBMRSC | C | 2/06 | 101.131 | 6.80 | 9.86 | 0.12 | 124.87 | 101.50 | 8.72 | 8.661 | 23.2 | 23.3 | 0.094 | 0.17 |  | 65.51 | 63.50 |  |  |  |  | - | - | - |
| BBMRSC | G | 3/05 | 100.551 | 6.97 | 9.97 | 0.56 | 126.32 | 100.30 | 8.65 | 8.661 | 22.6 | 22.6 | -0.007 | 0.08 |  | 65.67 | 73.50 |  |  |  |  | -0.09 | - | - |
| BBMRSC | J | 4/03 | 101.229 | 6.96 | 10.05 | 0.54 | 126.10 | 100.30 | 8.58 | 5.578 | 23.1 | 23.091 | 0.103 | 0.19 |  | 61.92 | 62.73 |  |  |  |  | -0.19 | 42.06 | 43.00 |
| BBMRSC | D | 5/06 | 103.621 | 6.94 | 10.01 | 1.75 | 126.25 | 100.7 | 8.54 | 8.498 | 23.567 | 23.6 | 0.062 | -0.09 |  | - | - | 0.00 |  | - | - | -1.13 | - | - |
| BBMRSC | B | 6/05 | 99.923 | 7.11 | 10.02 | 1.25 | 121.91 | 99.5 | 8.33 | 8.341 | 24.263 | 24.3 | 0.062 | -0.19 |  | 49.50 | 62.51 | -0.02 |  | 11.79 | 15.53 | -6.71 | 54.97 | 41.00 |
| BBMRSC | C | 7/01 | 101.097 | 7.17 | 10.11 | 3.99 | 119.80 | 104.30 | 7.759 | 8.10 | 24.560 | 24.6 | 0.061 | -1.08 |  | 55.35 | 61.41 | -0.18 |  | 13.40 | 15.20 | -2.36 | 34.06 | 41.00 |
| BBMRSC | A | 7/29 | 101.371 | 7.17 | 10.10 | 0.41 | 120.89 | 101.60 | 8.48 | 8.341 | 24.4 | 24.5 | -0.011 | 0.100 |  | 51.650 | 60.52 | 0.00 |  | 12.79 | 15.02 | -2.64 | 33.64 | 41.00 |
| BBMRSC | M | 9/02 | 100.590 | 7.13 | 9.99 | 0.31 | 118.20 | 100.80 | 8.38 | 8.341 | 24.617 | 24.7 | -0.023 | 0.05 |  | 51.47 | 61.96 | 0.00 |  | 12.68 | 15.38 | 0.03 | 33.50 | 41.00 |
| BBMRSC | E | 10/2 | 101.080 | 7.04 | 10.04 | 0.63 | 123.57 | 100.2 | 8.53 | 8.498 | 23.468 | 23.5 | - | 0.25 |  | 58.71 | 63.63 | 0.08 |  | 14.55 | 15.83 | 0.43 | 36.49 | 43.00 |
| BBMRSC | G | 11/4 | 0.997 | 7.11 | 10.03 | -0.24 | 124.71 | 100.7 | 8.60 | 8.578 | 23.300 | 23.2 | 0.083 | -0.01 |  | 62.31 | 64.00 | -0.02 |  | 15.47 | 15.92 | -0.04 | 41.09 | 41.00 |
| BBMRSC | M | 11/24 | 100.426 | 7.05 | 10.05 | 1.26 | 121.05 | 102.8 | 8.74 | 8.578 | 23.036 | 23.1 | 0.127 | 0.03 |  | 62.80 | 64.00 | 0.01 |  | 13.90 | 15.92 | -0.06 | 41.04 | 43.00 |
| BBMRSC | H | 12/23 | 100.433 | 6.97 | 9.97 | 0.29 | 124.49 | 100.8 | 8.72 | 8.661 | 22.475 | 22.5 | 0.092 | -0.08 |  | 63.16 | 63.75 | 0.00 |  | 15.65 | 15.86 | -0.63 | 42.41 | 43.00 |

14)  Other remarks/notes –

Data are missing due to equipment or associated specific probes not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform. If additional information on missing data is needed, contact the Aquatic Preserve office. Copies of the calibration/deployment logs can be obtained through the Water Quality Specialist.

All data files are QAQC’d using the 2012 CDMO Excel macro. This macro automatically flags DO values less than 3 mg/L as Passed Initial QAQC Checks with the comment DO Hypoxia (<3 mg/L) (0, CDA). The DO flag is based on Federal standards. During the primary QAQC session, any reading that falls outside a sensor’s range (see section II. Physical Structure descriptors) is rejected (-3). Additionally, negative turbidity readings between (0) and (-2) have been automatically flagged as suspect data (1, CAF) since the value is within the accuracy of the probe per the CDMO operations manual. All negative chlorophyll readings are labeled as (-4) since they fall outside the chlorophyll sensor range and are manually rejected during QAQC (-3, SNV). All the first data points for each new deployment are labeled (0, CND).

The parameters that are QC’d are depth (m), temperature (°C), salinity (ppt), specific conductance (mS), turbidity (FNU), pH (SU), DO (mg/L and %) and chlorophyll (ug/L). All other parameters have not been QC’d.

All data collected after the 30-days in each deployment is flagged suspect (1, GSM)

All turbidity readings over 124 FNU are flagged suspect (1, STS) since they are out of the calibration range that BBAP applies (0-124 FNU) but within acceptance range for YSI EXO2 datasondes as according to user manuals. All turbidity readings over 1000 FNU are 10x above calibration range and are rejected for all BBAP stations. All chlorophyll readings over the upper calibration value are flagged suspect (1, SCS) since they are out of the calibration range (varies by deployment). All negative Chl values are rejected (-3, SNV) Additionally, for BBAP, anomalous turbidity and chlorophyll spikes are labeled suspect (1, STS/SCS) if the previous and next data points are not in a defined “peak” or turbidity event and if the reading is more then STANDARD DEVIATION*3 then point is rejected (-3, STS, CSM) (SD calculated after rejected data removed) These readings may either be caused by optical interference by animals or fouling, or an unknown local disturbance. Readings within a well-defined peak were maintained and given a (0) flag.

For BBLR03, turbidity values > 12 FNU and Chlorophyll values > 6 ug/L from the previous and next data point were flagged suspect (1, STS/SCS).

For BBBB14, turbidity values > 7 FNU and Chlorophyll values > 3 ug/L from the previous and next data point were marked (1, STS/SCS).

For BBJT71, turbidity values > 5 FNU and Chlorophyll values > 3 ug/L from the previous and next data point were marked (1, STS/SCS).

For BBMRDW, turbidity values > 5 FNU and Chlorophyll values > 2 ug/L from the previous and next data point were marked (1, STS/SCS).

For BBMRSC, turbidity values > 12 FNU and Chlorophyll > 4 ug/L from the previous and next data point were marked (1, STS/SCS).

For BBMRRB, turbidity values > 8 FNU and Chlorophyll values > 4 ug/L from the previous and next data point were marked (1, SCS).

For BBCWA4, turbidity values > 10 FNU and Chlorophyll values > 4 ug/L from the previous and next data point were marked (1, STS/SCS).

January 1-31, 2025

BBLR03

Missing: Depth – 1/13 9:00 – 2/5 13:30;

Reject:

Suspect:; TB – 1/4 7:30 1/10 9:15 anomalous single point or group spike;

Notes: DO – 1/1 2:30-8:00 1/1 17:30-20:15 1/2 4:00-8:45 20:00-21:15 1/3 3:45-9:15 1/3 20:45-22:15  hypoxic conditions;

BBBB14

Missing: TAL – 1/23 12:30-2/20 9:30 sensor error no data recorded;

Reject: TAL – 1/1 00:00-1/23 12:00 negative values;

Suspect: TB – 1/5 6:45 1/9 11:30 1/11 5:45 1/17 5:15 5:45 1/18 6:00 1/21 8:30 1/22 9:00 9:45 10:30 1/28 18:00 1/30 14:00 1/31 12:45  anomalous single point spike;

Notes: DO – 1/18 6:45-8:15 1/20 9:30-10:15 1/21 00:30-2:45 1/22 2:30-2:45 3:45-4:00 4:45-5:00 1/23 00:45-1:30 1/20 8:30-12:15 20:30-20:45 23:45-1/21 1:00 10:15-13:00 21:15 1/22 1:15 1/22 8:00-14:15  hypoxic conditions; TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBJT71

Missing: Depth – 1/23 11:30-2/20 10:15;

Reject:

Suspect: All – 1/18 12:00-1/23 11:30

Notes: DO – 1/18 6:30-6:45 hypoxic conditions;

BBMRDW

Missing: All – 1/7 10:15;

Reject: TAL – 1/18 19:00–2/5 10:45 sensor consistently drifted upwards from 10ug/L to 80+ug/L and failed post cal.

Suspect: TB – anomalous single point spikes 1/11 18:00 1/19 13:00 1/26 14:45 1/27 9:30 1/31 23:15, TB – anomalous group turb spikes 1/13 12:30-14:00 1/14 15:15-16:15  1/19 14:15-30 1/29 16:15-17:45, 1/30 9:45-10:00,  ,

Notes: turb spikes were extreme, higher values and all grouped rather then normally seen single point spikes at this site.

BBMRSC

Missing:

Reject:

Suspect: ; TAL – 1/9 3:45 anomalous single point chlorophyl spike;

Notes: DO – 1/1 2:45-7:30 15:00-22:45 1/2 3:00-11:45 1/2 15:30-1/3 00:00 4:00-12:15 16:30-1/4 1:45 8:30-1/6 2:30 6:15 9:00-11:30 16:45-1/7 3:00 7:30-8:00 9:30-10:15 18:30-1/8 10:30 1/9 7:45-12:45 19:30-20:00 22:00-1/10 14:30 18:30 22:30-1/11 4:30 7:00-9:45 13:15 14:30 15:15 18:00 19:15-19:30 1/12 2:00-11:00 15:30-1/14 5:30 1/15 1:30-7:00 15:00-20:00 1/16 1:30-7:15 14:00-15:00 17:15-20:45 1/17 1:00-18:30 1/18 6:15 1/19 17:30-1/20 00:45 etc… hypoxic conditions; *consistently hypoxic and even anoxic conditions observed especially when salinity drops indicating releases from structure, this month saw salinity levels consistently bellow 5psu.*

BBCWA4

Missing:

Reject:

Suspect: All – 1/20 10:45–1/23 10:30 deployed past 30 days, data appears to fit conditions; Depth – 1/23 10:00-2/20 12:45 data appears to have recorded at slightly shallower then actual depth as seen in verification, improper calibration suspected to be the issue. SpCond – 1/1 00:00-1/23 9:45 C/T was reading slightly lower sal & sp.con then actual value as confirmed with verification and field sampling, see calibration log, sensor may have internal issue, flagged suspect; TAL – 1/1 00:00-1/23 9:45 passed Cal ICV low range CCV but failed high range CCV,

Notes:

BBMRRB

Missing:

Reject: TB - 1/30 16:15 single point turb spike far exceeding normal monitoring ranges;

Suspect: All – 1/20 10:45–1/23 10:30 deployed past 30 days; 1/23 16:00 sal/spcond single point drop; TB –1/8 11:00 1/16 22:00 1/24 16:30 1/27 9:00 anomalous single point turb spike;

Notes:

February 1-28, 2025

BBLR03

Missing: ALL - 2/5 13:45 new deployment began;

Reject: TAL – 2/7 2:00-15 11:45 2/14 11:30 2/16 9:00 2/18 10:45 11:30 12:45 2/25 7:45 12:15 2/26 8:00 9:45 11:30 2/28 9:30 11:15-30 negative value;

Suspect: All – past 30 day deployment 2/03 13:00–2/05 13:30; TB – 2/3 6:45-7:00 8:15 2/5 8:45-2/5 10:15 2/5 8:45 9:00-9:45 10:15 anomalous single point or group spike; TAL – 2/7 2:15-2:30 11:45 2/8 3:00 2/9 5:30

Notes: DO – 2/2 19:45 2/3 18:00 18:45 20:30-21:45 2/4 19:30-22:15 2/5 21:00-23:45 2/6 21:00-2/7 00:45 11:15-13:15 2/10 1:45-6:45 2/11 5:30-5:45 2/12 4:00-5:30 2/13 3:15-4:15 2/14 4:30-7:15 2/14 11:15 2/15 4:00-8:45 14:15-15:45 2/16 3:45-8:45 2/17 5:15-8:00 2/21 5:45-10:30 2/22 00:30-00:45  hypoxic conditions;

BBBB14

Missing: TAL - 1/23 12:30-2/20 9:30 sensor error no data recorded;

Reject:

Suspect: TB – 2/3 10:30 2/4 7:15-30 2/5 18:30 2/7 6:30 18:30 2/8 7:00 9:15 2/10 2:45 10:00 2/13 00:30 7:30-45 2/15 7:00 2/17 12:30 15:30 16:30 2/18 5:00  anomalous single point or group spike;

Notes: DO – 2/18 22:15-2/19 2:15 3:45-9:15 2/20 22:15-23:15 2/21 00:00-1:30  hypoxic conditions; TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBJT71

Missing: Depth – 1/23 11:30-2/20 10:15; All - 2/20 10:30;

Reject:

Suspect: TB – 2/15 9:30 anomalous single point turb spike;

Notes: DO – 2/19 2:00-2:30, 4:30, 5:00-7:30  hypoxic conditions

BBMRDW

Missing:

Reject: TB – 02/02/2025 13:45 anomalous single point turb spikes far exceeding normal monitoring range; TAL – 1/18 19:00–2/5 10:45 sensor consistently drifted upwards from 10ug/L to 80+ug/L and failed post cal

Suspect: TB – anomalous single point spikes 02/02/2025 22:15 02/03/2025 09:00 02/04/2025 11:30 02/04/2025 13:30 2/16 23:45, anomalous group turb spikes 02/03/2025 17:00-17:15, 02/04/2025 12:45-13:00

Notes:

BBMRSC

Missing: Depth – 2/5 13:00-2/9 12:15; All – 2/5 11:45-12:45 2/21 12:15

Reject: Depth – 2/21 12:45 anomalous single point depth increase; pH – 2/21 15:45-3/20

TB – 2/21 12:45 anomalous single point turb spike;

Suspect:

Notes: *consistently hypoxic and even anoxic conditions observed especially when salinity drops indicating releases from structure could be source of low DO events.*

BBCWA4

Missing: All – 2/20 13:00 2/23 10:00 2/24 00:15 7:00 17:45 19:00 2/25 12:30 2/28 1:15; TB – 2/24 7:15 anomalous single point or group spike over 3* standard, 3/4 11:45 13:00 3/6 8:30 anomalous single point or group spike and was suspect for additional reasons/flags; TAL – 2/23 5:00-3/20 23:45 sensor malfunction;

Reject: pH – 2/21 15:45-3/20 23:45 sensor malfunction; TAL – 2/13 4:30 anomalous single point or group spike;

Suspect: All – 2/23 4:00-3/20 23:45 instrument malfunction, telemetry data recovered, sonde data missing; Depth – 1/23 10:00-2/20 12:45 data appears to have recorded at slightly shallower then actual depth as seen in verification, improper calibration suspected to be the issue; TB – 2/17 13:30 2/21 8:30  anomalous single point or group spike; TAL – 2/3 20:00 2/4 14:30-15:00 2/7 22:00 anomalous single point or group spike;

Notes:

BBMRRB

Missing:

Reject: TB – 2/1 16:30 2/5 19:00  2/7 7:00 2/7 8:00 2/3 17:00 2/8 16:15 2/9 16:00 2/15 12:45 2/18 00:15 2/19 17:00 2/19 17:45 18:00 2/19 22:00 2/20 8:30 2/23 10:45 2/23 18:30  single point or group spike over 3*standard;

Suspect: All - 2/7 7:00 spcond single point drop; TB - 2/2 18:15 2/4 8:45 2/5 17:00 17:30 2/6 7:15 2/7 13:15 2/7 2/8 1:15 2/8 16:30 23:45 2/10 13:15 2/13 13:00 2/16 10:45 2/18 0:15 2/18 21:15 2/19 22:00 2/20 8:15 2/20 11:45 2/28 10:00  anomalous single point or group spike;

Notes:

March 1-31, 2025

BBLR03

Missing:

Reject: TAL – 3/1 5:45 9:45-10:15 11:15 3/2 8:30-8:45 10:30-11:00 3/3 8:30 11:45 3/5 9:30 10:30 negative value;

Suspect: TB – 3/5 14:30 3/6 9:30 3/20 12:15 anomalous single point and group spike;

TAL – 3/31 22:30 anomalous single point spike;

Notes: DO – 3/17 5:00-6:15 18:00-19:30 3/26 1:00-2:45 12:15-13:15 3/27 1:45-3:30 3/29 13:00-13:15 3/30 2:00-6:30 3/31 20:45 hypoxic conditions;

BBBB14

Missing:

Reject:

Suspect: TB – 3/1 20:30-45 3/2 18:30 3/4 8:45 3/5 17:00 3/6 8:30 3/21 11:00  3/26 9:15-30   anomalous single point spike; TAL – 3/1 20:30 3/26 9:30 anomalous single point spike;

Notes: TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBJT71

Missing: All – 3/20 9:30;

Reject:

Suspect: TAL – 3/20 9:45-3/26 10:15 negative chlorophyl value; TB – 3/19 16:15 random single turb spike;

Notes: decommissioned permanent monitoring on 3/26/2025 10:15

BBMRDW

Missing:

Reject: TB – anomalous single point spikes far exceeding monitoring norms 3/12 7:15

Suspect: TB – anomalous single point spikes  3/29 16:15 3/22 14:30

Notes:

BBMRSC

Missing:

Reject:

Suspect:

Notes; *consistently hypoxic and even anoxic conditions observed especially when salinity drops indicating releases from structure could be source.*

BBCWA4

Missing: All - 3/1 14:45 3/2 17:30 18:45 3/3 4:30 9:30 17:30 3/4 00:45-8:00  3/5 23:00 3/7 22:45 3/8 9:00 3/9 14:45 22:30 3/13 2:15 3/14 00:45-8:00 3/15 00:00 ; TAL – 2/23 5:00-3/20 23:45 sensor malfunction no data collected 3/31 2:00 anomalous single point or group spike over standard;

Reject: pH – 2/21 15:45-3/20 12:15 sensor malfunction; TAL – 3/31 2:00 anomalous single point Chl-a spike,

Suspect: All – 2/23 4:00-3/20 23:45 instrument malfunction, telemetry data recovered, sonde data missing TB – 3/26 6:30 3/28 8:30  anomalous single point turb spike,

Notes:

BBMRRB

Missing:

Reject: TB - 3/16 6:45 3/16 7:15 3/16 18:00 single point turb spike far exceeding normal monitoring ranges;

Suspect: TB – 3/4 9:45 3/9 18:45 3/11 4:00 3/16 9:00 3/23 11:15 3/29 4:30 anomalous single point turb spikes;

Notes:

April 1-30, 2025

BBLR03

Missing:

Reject: TB – 4/17 00:30 anomalous single point or group spike over 3*standard; TAL – 4/18 9:00 4/28 15:15 anomalous single point spike over standard;

Suspect: TB – 4/4 15:30 4/7 7:30 4/18 11:45 anomalous single point or group spike,

Notes:

BBBB14

Missing:

Reject:

Suspect: TB – 4/6 22:15 4/7 14:30-45 4/7 18:00 anomalous single point spike, TAL - 4/20 23:45 anomalous single point spike;

Notes: TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBMRDW

Missing:

Reject:

Suspect: TB – 4/5 20:45 4/12 14:15 4/5 13:45 4/21 19:45 4/18 19:45 4/22 22:45 4/24 14:00 4/29 18:15 4/30 21:00 anomalous single point turb spike, 4/1 22:15-30 4/27 22:15 22:45 anomalous group turb spike TAL – failed posted cal all suspect

Notes:

BBMRSC

Missing:

Reject:

Suspect: TB – failed post call all suspect but 4/2 7:00 5/23 18:15 anomalous single point turb spike; TAL – all failed post cal

Notes:

BBCWA4

Missing:

Reject: TB – 4/2 00:45 4/6 1:00 4/8 3:00-15 2:00 4/12 2:00 3:30 4:00 19:15-30 21:00 21:45 22:00 23:15 23:45 4/13 2:00 2:30 3:30 4:45 5:15 8:30 13:15 19:45 20:45 23:30 4/14 21:15 22:15 23:45 4/15 5:00 4/29 16:45 4/6 2:00 4/10 20:15-21:30 4/11 3:15 3:45 20:30-22:15 4/12 20:00-30 22:30-45 4/13 21:45-22:15 4/14 19:15-45 22:45-23:00 4/15 4:15-30 4/29 16:45  anomalous single point or group spike over 3*standard; TAL – 4/3 22:15 4/20 10:30 anomalous single point spike above standard;

Suspect: C/T (All)– 4/5 10:00 4/12 23:45 4/13 2:30-2:45 4/13 13:15 4/13 23:15-23:30 4/14 23:45 unexplained cond. drop; TB – 4/2 21:45 4/3 21:00 4/5 9:30 10:00 4/7 19:45 20:30 4/9 20:30  5:30 7:00 4/10 1:00 3:15-30 4:00-15  5:30 7:00 19:30-20:00 21:45 22:30 23:00  4/11 5:00 4/11 5:30 4/12 2:00, 4/6 1:15-1:30 2:30-45 3:15-30 4/7 1:00-30 2:15-3:00 3:30-45 4:30-5:15 20:30 4/8 2:00-5:30 4:15 4/10 1:00 2:15-45 3:15-30 4:00 19:30-21:30 4/11 4-4:30 5:00 5:30 19:30-23:00 4/12 3:45-5:30 19:15-4/13 5:30 8:30 13:15 19:45-23:30 4/14 20:15-4/15 00:45 4:45 5:15-5:30, 4/15 19:30-45 20:15-45 22:00-23:00 4/16 00:15-45 1:15 2:30-3:15 4:00 4:45 4/17 00:15 1:15 5:30 20:30 4/18 00:45 1:15-45 2:15 2:45 3:45 5:30 4/19 00:45 2:00 3:00 3:30 3:45 4:30 5:30 20:15 21:15 4/20 2:15 2:45 3:45 4/21 00:15-45 2:00 2:45 22:15 22:45 4/22 5:00 23:30 4/24 17:45 4/25 7:30 19:15 4/26 3:15 3:45-4:00 20:00 4/27 17:15 4/28 5:15 21:15 4/29 22:00 4/30 13:15 anomalous single point or group spike; TAL – 4/3 20:00 20:30 4/5 15:30 4/11 23:30 4/23 8:30 4/30 20:45 anomalous single point spike,

Notes:

BBMRRB

Missing:

Reject: TB – 4/5 14:45 4/7 10:45 12:30 4/15 21:15 anomalous single point turb spike

Suspect: TB – 4/7 12:45 15:15 16:00  anomalous single point turb spike, anomalous group turb spike

Notes:

May, 2025

BBLR03

Missing:

Reject:

Suspect: All – 5/3 11:45-5/6 11:45 Instrument deployed past 30 days, TB – 5/7 4:00 5/16 15:45 5/17 14:15 5/25 15:15 anomalous single point or group spike,

Notes:

BBBB14

Missing:

Reject:

Suspect: TB – 5/13 5:30 5/29 2:30 anomalous single point spike, TAL – 5/9 9:45-6/13 9:00 failed post calibration 5/19 6:15 5/29 2:30 anomalous single point spike,

Notes: TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBMRDW

Missing:

Reject:

Suspect: All – 5/3 9:45-5/6 9:45 instrument deployed past 30 days TB – 5/1 17:30 5/2 17:30 5/3 18:00 20:15 5/4 4:30 10:15 16:00 5/6 19:45 5/10 21:15 5/11 15:15 5/16 10:30 5/19 9:45 5/21 15:45 anomalous single point turb spike, anomalous group turb spike TAL – failed posted cal all suspect

Notes:

BBMRSC

Missing: All – 5/6 10:30 New deployment began

Reject:

Suspect: TB – failed post call all suspect TAL – all failed post cal

Notes:

BBCWA4

Missing: All – 5/9 12:45 New deployment began, TB – 5/2 15:15 5/9 6:30 5/10 3:15

Reject: TB – 5/2 15:15 5/9 6:30 5/10 4:00 20:30 5/17 9:30 10:00 5/20 11:30 5/23 11:15 5/28 8:45 15:45 anomalous single point spike, TAL – 5/9 14:00 5/17 7:15  anomalous spike and failed post calibration,

Suspect: TB – 5/1 20:30 22:00 5/2 2:15 19:30-45 22:00-15 23:00 5/3 1:00 1:45 5/5 00:45 2:30 22:45 5/8 3:15 5/10 4:15 4:45 21:15 5/15 00:15 6:45 17:30 5/16 6:00 12:15 5/17 9:15 13:30 20:30 21:00 23:30 5/18 5:30-45 11:45 5/19 2:45 10:00 11:15 16:00 5/20 13:15 5/24 16:00 18:00 5/25 14:00 5/28 16:45 - anomalous single point spike or group spike; TAL – 5/9 13:00-6/13 10:30 5/27 14:45 post calibration out of range,

Notes:

BBMRRB

Missing: All – 5/9 11:45 new deployment began;

Reject: TB – 5/5 23:30 5/6 23:45 5/8 00:15 22:15 5/11 5:45 5/13 22:15 5/25 17:30 anomalous single point turb spike, TAL – 5/11 12:30 13:00 13:30 negative value,

Suspect: TB –5/7 19:30 5/9 4:45 5/14 14:45 18:15 5/20 6:00 5/20 20:00 5/21 4:00 14:31 5/22 21:30 5/24 18:45 5/25 15:15 5/27 10:15 5/31 11:30 anomalous single point turb spike, 5/9 12:15-6/2 13:30 post calibration out of range but data appears to fit conditions; TAL – 5/9 12:00-6/13 11:15 post calibration out of range

Notes:

June, 2025

BBLR03

Missing:

Reject: Sp. Cond & Sal – 6/5 12:15-7/1 10:15 passed Cal. and ICV but appears off once deployed as initial deployment verification showed similar values to retrieved sonde but once deployed saw sal at 10psu and sp.cond at 20 ms/cm above actual reading this was shown with next deployment verification on 7/1

Suspect: TB – 6/9 15:45 anomalous single point or group spikes; TAL – 6/5 12:15-7/1 10:00 passed Cal. ICV and low range CCV but failed high range CCV, 6/24 10:30 anomalous single point spike,

Notes:

BBBB14

Missing: All – 6/13 8:45 New deployment began,

Reject:

Suspect: All – 6/9 9:45-6/13 9:00 Instrument deployed past 30 days; TB – 6/12 6/16 13:30 15:15 anomalous single point or group spike; TAL – 5/9 9:45-6/13 9:00 passed Cal. ICV low range CCV and failed high range CCV, 6/1 17:30 6/12 15:30 6/13 9:00 6/19 10:30 6/22 7:30 6/23 7:00-30 6/30 14:30 anomalous single point or group spike,

Notes: TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBMRDW

Missing:

Reject:

Suspect: TB – 6/1 20:30 6/7 14:15 6/12 21:30 6/15 12:45 6/16 21:45 6/24 3:15 6/25 7:45 6/29 6:45 8:30 6/21 17:00-15 6/28 20:45-21:00 anomalous single point or group turb spike

Notes:

BBMRSC

Missing:

Reject:

Suspect: TB – failed post call all suspect TAL – all failed post cal

Notes:

BBCWA4

Missing:

Reject: TB – 6/3 10:30 6/5 8:45 10:00 6/6 19:30 6/8 13:45 16:30 18:00 20:00 6/9 12:00 12:30 6/10 13:00 22:45 6/13 12:00 6/21 5:30 6/23 17:30 6/26 7:45 16:15 6/27 5:30 13:00-15 14:45 18:15 19:30 6/28 8:15 10:00 11:15 15:30 18:30 6/30 8:45 anomalous single point spike,  TAL – 6/6 00:30 6/7 21:15 6/21 14:30 6/28 10:00 anomalous single point spike and failed post calibration,

Suspect: ALL – 6/9 13:00-6/13 10:45 deployed past 30 days but conditions appear to fit recorded data; C/T – 6/16 19:30-7/9 8:30 suspect due to diagnostics and prong appeared to have small cracks along epoxy of sensor; TB – 6/5 10:30 6/7 1:45 2:30-45 6/9 10:30 6/14 11:15 20:30 6/15 3:15 7:45 11:15 13:15-30 14:45 6/16 13:45 18:30 19:15 6/17 6:00 14:00 6/18 5:45 14:15 6/19 10:30 18:15 18:45 6/20 00:45 6:00 10:15 11:00 6/21 1:15 5:00 5:30 6:00 6/22 10:00 12:30 6/23 7:15 8:15 17:15 6/24 12:00 15:00 6/28 7:15 10:00 6/29 16:30 6/30 18:00  - anomalous single point or group spike; TAL – 6/26 13:00-15  anomalous single point or group spike;

Notes:

BBMRRB

Missing: All – 6/2 2:30 13:45-6/13 11:30 unsure why data is missing

Reject: TB - 6/14 18:45 6/15 10:15 6/18 14:45 6/21 19:156/27 11:00 anomalous single point spike; TAL – 5/25 13:00-6/2 13:45 sensor drift and failed post calibration,

Suspect: All -  6/9 12:45-6/13 10:45 Instrument deployed past 30 days TB –6/14 7:30 10:00 6/14 18:00 6/16 10:30 6/17 19:30 6/18 8:15  6/19 16:00 6/21 7:30 18:45 6/23 3:00 5:30 10:15 12:00 6/24 16:00 6/26 19:00 6/27 00:00 19:15 6/28 15:15 6/29 12:00 14:15 6/30 14:00 7/2 19:15 7/40:15 12:00 6/24 16:00 6/26 19:00 6/27 00:00 11:00 19:15 6/28 15:15 6/29 12:00 14:15 6/30 14:00 7/2 19:15 7/4 6:45 7/5 11:15 14:30 anomalous single point turb spike,

Notes:

July 1-31, 2025

BBLR03

Missing:

Reject: TB – 7/29 12:45 anomalous single point or group spike over 3*standard; TAL – 7/1 11:15 negative value;

Suspect: DO – 7/1 10:30-7/29 11:00 sensor passed Cal & ICV but failed CCV possibly due to biofouling from small calcareous worm tubes on sensor tip;

Notes: Regular instances of hypoxia

BBBB14

Missing:

Reject:

Suspect: TB – 7/6 10:00 7/10 19:00 anomalous single point spike, TAL – 7/09 10:45-8/08 8:45 passed Cal ICV low range CCV failed high range CCV;

Notes: TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBMRDW

Missing:

Reject: TB – 7/20 10:00-10:45 7/23 11:00 anomalous group spike 7/21-7/29; Chl – 7/25 9:45-7/29 9:15 sensor drifting upwards towards 20 ug/L but verification with next sonde values only recorded 4 ug/L, unsure as to cause of sensor drift but high Chl levels recorded most of this deployment;

Suspect: TB & Chl – 7/01 9:15-7/29 10:30 passed Cal and ICV but failed both CCV’s possibly due to sensor drift;

Notes: Regular instances of hypoxia

BBMRSC

Missing: 7/29 8:15 new deployment begins;

Reject:

Suspect: DO, TB & Chl – 7/01 9:00-7/29 8:00 passed Cal and ICV but failed CCV;

Notes: Regular instances of hypoxia

BBCWA4

Missing:

Reject: TB – 7/2 5:45-6:00 10:30 13:30 7/6 9:15 7/12 8:15  7/13 9:30 7/17 9:30 10:45 11:30 12:45 7/18 7:45 12:30 14:30 15:15 19:30 7/19 9:15 16:45 17:00 7/20 5:30 7:30 8:45 7/21 6:30 9:45 15:30 7/24 10:45 11:45 20:00 7/25 8:45 10:00 18:00 7/19 20:15-21:30 7/20 14:45-15:00 7/22 8:15-9:30 17:45-18:00 19:00-19:15 7/24 16:45-17:15 7/26 9:15 10:30 13:30 14:30 7/27 8:15 12:30 17:30 19:15-7/28 5:45 10:15 12:00 19:30 23:15-30 7/29 2:15 8:15 8:45 11:00 7/31 5:30 11:15 12:30 anomalous single point or group spike;

Suspect: C/T – 6/16 19:30-7/9 8:30 suspect due to diagnostics and prong appeared to have small cracks along epoxy of sensor; TB – 7/1 18:45 7/2 11:15 18:15 19:00 7/4 7:45 7/5 8:30 9:45 10:45 12:30 19:15 7/7 18:45 7/8 18:15  anomalous single point or group spike, TAL – 7/1 5:00-15 anomalous single point spike,

Notes:

BBMRRB

Missing:

Reject: TB – 7/6 18:15 19:15 7/9 6:45 7/19 11:45 7/20 6:00 10:15 15:00 17:00 7/19 11:45 7/21 16:30 7/22 13:15 7/23 10:45 7/24 7:30 17:00 7/25 6:45 10:30 14:00 7/28 18:30 20:45 7/29 7:00 16:00 16:30 anomalous single point spike, 7/10 5:30-6:00 7/18 17:45 7/22 7:30 8:30 7/25 16:30-45 7/26 10:15 11:30 12:00-15 16:45-17:00 7/27 6:45 7/28 21:00 7/29 7/30 5:00 11:15 7/31 13:00-15 anomalous group spike; TAL -

Suspect: TB – 7/9 18:30 7/17 23:00 7/23 6:45 anomalous single point spike, 7/10 5:30-5:45 7/15 13:15-14:30 20:45 7/18 9:15 11:00 12:45 17:00 7/19 14:00 7/20 9:00 14:00 7/21 9:15  7/22 300 8:45 7/26 10:30 11:45 7/27 6:30 13:30 14:30 7/28 00:30 11:45 anomalous group spike; TAL – 7/9 9:15-; TAL – all passed Cal and ICV but failed CCV

Notes:

August 1-31, 2025

BBLR03

Missing: All – 9/2 10:45 new deployment began;

Reject: TB -  anomalous single point spike; TAL - 8/21 18:30 anomalous single point spike over 3*standard;

Suspect: TB - 8/2 14:30 8/8 10:45 8/16 00:15-30 8/21 7:45 18:30 anomalous single point spike; TAL – 8/6 18:00 8/13 14:00 15:00 23:30 8/14 14:30-14:45 15:45 8/21 18:30 anomalous single point spike;

Notes: Regular instances of hypoxia

BBBB14

Missing: All – 8/8 9:00 new deployment began

Reject:

Suspect: TB – 8/8 9:15-9/9 10:45 verification & graph appears to be recorded 1FNU higher then actual conditions; TAL – 8/2 10:00 anomalous single point spike, 8/8 9:15-9/9 11:00 passed Cal ICV low range CCV failed high range CCV;

Notes: TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBMRDW

Missing:

Reject:

Suspect:

Notes: Regular instances of hypoxia

BBMRSC

Missing: All – 7/29 8:15 new deployment began; Depth – 7/29 8:30-9/2 8:15 sensor dropped out, no data collected;

Reject: TB – 8/7 3:45 anomalous single point spike;

Suspect:

Notes: Regular instances of hypoxia ranging into anoxic conditions

BBCWA4

Missing: All – 8/8 10:15 new deployment began;

Reject: TB – 8/1 8:00 8/2 10:45 11:30 8/3 8:45 14:45 8/4 8:30 9:00 9:45-10:00 10:45 8/4 15:00-16:15 8/5 2:15 9:00-9:15 8/6 2:45 6:00 8:00 9:45 8/7 5:45 14:00 16:15 18:15 8/8 19:00 8/11 6:15 8/27 8:30 8/28 14:45 anomalous single point spike;

Suspect: TB – 8/2 9:15-30 8/4 7:15-30 17:15 18:30 8/5 6:45 8/7 1:15 anomalous single point spike, 8/7 12:15-13:30 8/8 7:45 anomalous group spike; TAL – 8/8 10:30-9/9 8:45 passed Cal and ICV but failed ICV;

Notes:

BBMRRB

Missing:

Reject: TB – 8/1 5:30 13:00 8/2 11:00 12:30 8/3 7:00 8/4 10:15 8/5 11:00 13:45 15:45 16:30 8/6 16:15 8/7 6:15 8/7 10:00 8/7 15:30 8/12 16:30 8/14 1:30 8/16 1:30 8/17 5:45 8/20 17:30 8/21 16:15 8/22 6:00 10:15 8/23 14:00 17:15 8/24 6:45 8:45 8/26 16:00 8/28 12:30 8/29 8:30 8/30 13:45 8/31 12:15 anomalous single point spike, 8/2 14:00-15 anomalous group spike;

Suspect: TB – 8/3 8:30 8/5 11:00 8/6 14:30 8/7 18:45 8/8 9:15 8/16 3:00 8/24 10:00 8/29 7:45 9:15 10:30  anomalous single point spike, anomalous group spike; TAL – all passed Cal and ICV but failed CCV

Notes:

September 1-31, 2025

BBLR03

Missing: ALL - 9/2 10:45 new deployment began;

Reject:

Suspect: TB – 9/14 16:00 9/21 8:00 9/26 9:00 anomalous single point spike; TAL – 9/2 11:00-10/2 10:00 passed Cal and ICV but failed CCV;

Notes: Regular instances of hypoxia

BBBB14

Missing: All – 9/9 11:00 new deployment began;

Reject:

Suspect: TB – 9/24 1:00 9/30 17:30 anomalous single point spike; TAL – 9/9 11:00- 10/9 8:30 passed Cal ICV low range CCV but failed high range CCV, 9/12 6:30 9/23 6:45-8:00 anomalous single point or group spike;

Notes: TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBMRDW

Missing: All - 9/16 8:00-10/2 9:00 SD card failure, possibly due to failure of cond. side of C/T sensor, no data recorded;

Reject: C/T – 9/2 9:45-9/16 8:00 conductivity sensor failure, temp was unaffected; Turb – 9/12 17:15 anomalous single point spike; TAL – 9/2 9:45-9/16 8:00 passed Cal and ICV but failed CCV;

Suspect:

Notes: Regular instances of hypoxia

BBMRSC

Missing: All – 9/2 8:30 new deployment began;

Reject: TB – 9/26 15:00 anomalous single point spike; Chl – 9/11 16:45 9/22 18:15 9/23 4:45 anomalous single point spike;

Suspect: TAL – 9/2 8:45-10/2 11:00 passed Cal and ICV but failed CCV;

Notes: Regular instances of hypoxia ranging into anoxic conditions.

BBCWA4

Missing: All – 9/21 22:00-9/22 2:15 9/22 22:00-9/23 2:30 18:30 22:00-9/24 2:00 8:00 9:30 22:00-9/25 2:15 22:00-9/26 2:30 9:45 18:15 22:00-9/27 2:15 22:00-9/28 2:00 9:45 12:30 22:00-9/29 2:45 9/29 22:00 21:45-10-01 2:15 instrument malfunction due to telemetry and battery issues;

Reject: TB – 9/5 9:15 9/7 6:15 anomalous single point spike, 9/4 18:30 9/8 6:15 8:00 8:45 16:30 9/17 16:15 9/19 12:30 9/20 8:30 14:45  anomalous group spike;

Suspect: Depth – 9/9 9:00-10/9 10:00 possible offset issue with calibration but data appears to be off; TB – 9/9 9:00-10/9 8:30 passed Cal and ICV but failed CCV, 9/1 15:00 17:15 9/2 18:30 9/3 6:15  anomalous single point spike, 9/1 16:30-17:30 9/4 18:15 9/5 5:45 14:15 17:00 9/6 6:30 9/6 14:45 9/7 6:45 16:45 9/9 6:15 9/20 9:45 9/21 10:00 anomalous group spike; TAL – all passed Cal and ICV but failed CCV

Notes:

BBMRRB

Missing: All – 9/9 8:45 new deployment began;

Reject: TB - 9/1 11:45 9/2 7:00 9/4 15:45 9/5 17:00 9/6 6:00 9/7 9:45 9/8 16:45 anomalous single point spike,

Suspect: TB – 9/2 18:30 9/12 5:45 9/18 15:00 9/21 11:00 9/26 3:30 5:30 9/30 8:45 anomalous single point spike; TAL – all passed Cal and ICV but failed CCV

Notes:

October 1-31, 2025

BBLR03

Missing:

Reject: TB – 10/23 15:45 10/25 9:00 anomalous single point spike over 3*standard;

Suspect: TB – 10/16 11:00 10/21 15:45 10/25 18:45 10/27 17:4510/28 21:30 10/30 9:00  anomalous single point spike;

Notes: Hypoxic conditions commonly recorded

BBBB14

Missing:

Reject: TB – 10/2 16:15 10/4 13:30 17:30 10/6 16:45 10/7 17:15 10/28 2:45 10/31 11:15 14:15 anomalous single point spike; TAL – 11/3 13:15-11/10 12:45 11/13 11:15  passed Cal ICV & CCV but sensor drifted values outside cal. range, 11/25 22:00 anomalous single point spike and failed CCV.

Suspect: TB – 10/1 8:00 12:45 15:15 10/2 13:15 10/3 13:30 22:45 10/4 9:30 13:00-15 15:00 17:45 10/5 00:15 12:15 16:15 10/7 12:00 15:30 16:00 10/8 13:00 17:30 10/21 2:45 10/23 21:00-15 22:00 10/24 2:00 2:30 10/24 20:00 20:45 22:30-45 10/25 1:45 3:30 5:30-6:00 10/25 189:30 19:15 10/25 23:30 10/26 7:00-30 10:15 12:30 10/27 00:00 10/27 11:15 10/27 17:30 20:30-45 21:45 23:30 10/28 4:30-45 15:15 18:45 19:45 10/29 4:45 10/29 14:45 16:15 17:30 19:15 21:00 10/30 10:00 12:00 13:45 17:15-30 21:00 21:45 10/30 21:00 21:45 23:15 10/31 00:00 1:00 3:00 3:30 4:15 4:45 5:45-6:15 7:00 7:45 8:15 8:45 9:45 11:30 12:45-13:00 14:15 16:45 20:45-21:00 22:15 anomalous single point spike; TAL – 10/09 8:45-11/3 13:15 11/10 12:45-11/13 11:15 passed Cal ICV & CCV but sensor drifted; 11/13 11:45-11/25 21:45

Notes: TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBMRDW

Missing:

Reject: TB – 10/27 9:00 anomalous single point spike more than 3* standard;

Suspect: ALL – 11/2 9:30-11/4 10:00 deployed past 30 days; DO – 10/2 9:15-11/4 10:00 passed Cal ICV but failed CCV by thin margin;

Notes:

BBMRSC

Missing: Depth – 10/2 11:45-11/4 9:00 no data recorded despite passing Cal ICV and CCV, possibly motherboard or specific sensor failure;

Reject:

Suspect:

Notes:

BBCWA4

Missing: All – 10/1 7:30 11:45 22:00-10/2 2:15 22:00-10/3 2:30 22:00-10/4 2:00 13:00 22:00-10/5 2:15 22:00-10/6 2:15 4:30 22:00-10/7 2:00 22:00-10/8 2:30 11:30 22:00-10/9 2:30 instrument malfunction due to telemetry or battery issues;

Reject: C/T, DO, pH – 10/09 17:00-11/2 8:00 Conductivity sensor failure all sensors reliant on C/T rejected, epoxy circles covering anodes on internal “U” side of wipe C/T sensor fell off while deployed causing exposure to elements and therefore inaccurate readings. TAL - 10/19 12:30 anomalous single point spike; TB – 10/24 10:15 anomalous single point spike;

Suspect: Depth, TB, TAL - 10/09 17:00-11/2 8:00 instrument malfunction caused by conductivity sensor failure, all data not reliant on C/T labeled suspect due to general instrument malfunction.

Notes:

BBMRRB

Missing: ALL – 10/22 6:15-11/13 12:30 battery failure no data recorded.

Reject: TB – 10/09 18:15 10/11 14:00 10/14 16:15 10/15 15:00 10/20 17:15 10/20 23:45 anomalous single point spike;

Suspect: TB –  10/9 18:00 10/10 2:30 10/11 3:15 10/14 8:15 10:00 10/18 17:30 10/20 14-14:15 18:00 19:00 10/21 8:45 14:45   anomalous single point or group spike; TAL –  10/11 15:45 10/12 4:00 10/13 15:15 anomalous single point spike, 11/13 13:00-12/10 16:15 passed Cal ICV low range CCV but failed high range CCV;

Notes:

November 1-30, 2025

BBLR03

Missing:

Reject: TB – 11/2 17:00 11/17 1:15-30 3*standard anomalous or group spike;

Suspect: TB – 11/1 11:00 11/19 15:15 11/20 22:45 11/24 9:45-10:00 11/26 14:00 11/26 15:45 anomalous single point or group spike; TAL – 11/4 11:00-11/25 11:00 passed Cal ICV & low end CCV but failed high end CCV

Notes: Hypoxic conditions commonly recorded

BBBB14

Missing:

Reject: TB - 11/3 13:00 11/3 19:15 11/4 14:00 11/9 5:45 11/10 17:45 11/13 10:45 11/16 4:45 11/24 1:15 11/25 00:00 2:15 3:15 4:15 18:30 20:30 11/26 19:15 11/27 2:45 3:30 20:30-45 11/30 19:00 20:45  anomalous single point spike over 3*standard; TAL – 11/17 14:30 11/25 22anomalous single point spike;

Suspect: TB – 11/1 1:00 1:30 3:00 3:45 4:15 5:30 6:00 7:15 8:15 8:45 11:00-15 12:30 13:30 18:15 20:00 21:00-45 22:45 23:30 11/2 00:15 3:30-4:15 4:45-5:15 8:30 11:30 13:00 16:00 17:15 21:00-15 21:45 11/3 00:30 1:00 1:30 2:00-3:00 5:00 8:00 10:15-30 11:15 13:15 13:45 - 14:45 16:30 18:30 11/4 00:00 1:30 2:45 3:15-30 4:15-5:00 6:00 7:00 8:15 11:30 13:30 16:00 18:15-30 19:00 20:00 20:30-45 21:15 22:45 11/5 00:30-45 2:00 2:30 3:00-45 11/5 5:15 7:00 7:30 8:15 10:00 10:45 12:30 14:30 11/5 19:15 19:45-20:00 21:15 11/6 1:30 2:15-30 4:45 5:45 8:30 10:45 13:15-30 14:45-15:15 16:15 17:30-18:00 18:30 19:00-19:30 20:45 22:15 11/7 1:00-15 4:00 4:45 5:45 7:15 7:45 9:00 10:00 12:45 16:15 17:00 19:00 21:00 21:30 11/8 2:15 3:15 6:15 11:45 12:45 13:45 17:45 18:15 19:15 20:00 21:00-15 22:00 11/9 1:45-2:00 2:45 3:15 5:00 6:15 6:45 7:45 8:30 9:30 13:15 14:30 15:45-16:00 17:30 19:00 19:30 20:30 22:15 22:45 11/10 1:30-45 3:00-15 4:45 9:00-15 11:00 11/10 11:45 12:30 13:00 15:30 17:15 18:00-15 22:15 11/11 00:15 3:45 4:45 5:45 6:15 10:00 13:45 14:15 14:45 17:15 19:30 20:30 21:00 21:45 22:15-23:00 23:30 11/12 00:00 00:45 2:15 3:00 5:00 11:30 12:15 13:00-15 14:15 15:45 16:15 17:15 18:00 18:30 20:00-15 20:45-21:00 21:30 22:15 23:00 23:30-45 11/13 00:45 1:45 2:30 3:45 5:15 6:15 7:15 10:00 11/13 16:30 20:15 11/14 12:00 23:00 11/15 3:30 15:00 19:45 23:45 11/16 6:45 15:30 16:30 22:45 11/17 2:15 17:00 21:30 22:15 11/18 00:00 6:15 7:30 15:15 20:15 22:45 11/19 1:45 2:45 3:45 11/19 10:30 17:30 11/20 00:15 1:00 2:45 7:30 8:15 19:30 23:15 11/21 1:15 2:15 10:30 11/21 23:45 11/22 00:30 5:45 14:15 21:00 11/23 1:15 4:15 4:30 10:45 18:4520:15 20:45 22:45 11/24 2:30 3:45 4:30 5:30 10:30 12:15 20:30 21:00 21:45-22:00 22:45-23:00 11/25 00:15 1:30 2:00 3:00 3:30-45 5:15 12:45 18:45-19:00 19:30 20:00-22:15 23:15-23:45 11/26 00:15-30 1:45 3:15 4:45 6:15 8:45 18:30-45 20:00 20:45 21:15-45 22:15-30 23:00-23:30 11/27 00:30-1:30 3:00 4:00 4:45 7:15 16:15 19:45 20:15 21:00 21:45 22:15 22:45 11/28 00:00 1:45 2:45 5:45 15:00 18:30 20:00 20:30 21:00 23:45 11/29 1:00-30 3:45 5:45 15:30 23:45 11/30 1:00 2:30 4:00 11:30 13:45 17:45 anomalous single point or group spike; TAL – 10/09 8:45-11/3 13:15 11/10 12:45-11/13 11:15 passed Cal ICV & CCV but sensor drifted; 11/13 11:45-11/25 21:45 11/25 22:15-12/11 1:15 passed Cal ICV low range CCV but failed high range CCV

Notes: TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBMRSC

Missing:

Reject: ALL – 11/12 20:45-21:15 possible tampering with device or mount; TAL - 11/8 11:15 11/11 10:15 10:45 11:30 12:00 11/12 0:00 00:45 2:00 3:00-15 6:30 10:15-45 11:15-12:15 12:45-13:15 13:45-15:45 16:45 17:30 11/13 00:45-4:15 5:00 5:30-45 6:15 7:00 8:30-10:30 11:30-13:45 14:15-17:00 11/14 2:30-3:30 4:15-45 5:15-30 6:00-15 6:45-7:45 8:15 8:45-9:00 9:30 10:00 10:30-12:00 12:30-13:15 13:45 14:30-45 15:15 16:00 16:30 11/15 8:30 10:15-11:00 11:45 12:15 11/16 8:00 11/17 7:30 8:15 8:45 9:30-11:45 11/18 2:45 3:30 5:30 6:45-:7:45 8:15 9:00-12:00 12:30-13:00 13:30 11/19 7:15 9:45-11:00 11:45 12:15 11/20 6:30 7:30-9:15 9:45-12:30 11/20 13:15 21:15 11/21 7:00-9:00 9:45-10:15 10:45-11:00 11:30-13:15 13:45-14:00 15:00 18:45 19:30 11/22 00:30-1:15 10:30-13:30 14:45 11/23 9:15-14:00 11/24 00:15 1:00 2:30-45 4:15 5:45 6:15 7:30 8:00 9:00-13:00 14:00 14:30 15:00 21:45 11/25 00:15-45 1:15 3:00-4:15 7:15-9:15 negative value;

Suspect: TB – 11/25 9:15-12/23 9:30 passed Cal ICV low range CCV but failed high range CCV; TAL – 11/4 9:15-11/25 9:15 passed Cal ICV CCV but data appears off by 1 FNU;

Notes:

BBMRDW

Missing:

Reject:

Suspect: TB – 11/30 23:15 anomalous single point spike;

Notes:

BBCWA4

Missing:

Reject:

Suspect:

Notes:

BBMRRB

Missing: ALL – 10/22 6:15-11/13 12:30 battery failure no data recorded.

Reject:

Suspect: TB –  TAL – 11/13 13:00-12/10 16:15 passed Cal ICV low range CCV but failed high range CCV,

Notes:

December 1-31, 2025

BBLR03

Missing:

Reject:

Suspect: TB – 12/2 13:15 12/25 14:15 anomalous single point or group spike;

Notes: Hypoxic conditions commonly recorded

BBBB14

Missing:

Reject: TB – 12/1 23:00 12/2 2:30 12/3 1:15 2:00 21:45 23:15 12/4 1:45 12/5 6:00 12/6 21:45 12/8 1:00 12/9 00:15 4:00 4:30 18:30 19:30 20:15 22:15 23:30-45 12/10 2:30 3:45 19:15 21:15 12/11 1:45-2:00 2:45 12/16 23:30 12/17 15:15 12/18 2:45 15:30 12/21 3:00 12/21 20:30 12/26 13:00 12/27 16:45 12/30 13:15 12/31 11:30  anomalous single point or group spike over 3*standard; TAL – 12/11 11:30-12:30 13:15-14:15 14:45-17:00 22:15-12/12 00:15 12-12 00:45-4:45 5:15-30 6:00-15 7:30-45 8:15-12:45 13:15-17:30 12/12 18:00-18:30 19:15-12/14 23:45 12/15 00:15-8:45 9:15 11:45-12:30 13:00-15 13:45-14:45 15:15-16:15 16:45-17:15 17:45 18:15-30 12/16 00:15-30 1:00-2:00 2:30-12/19 20:00 20:30-12/23 10:30 12/23 11:00-12/25 3:15 12/25 4:00-17:15 17:45-12/30 20:30 12/31 00:00-15 00:45-21:00    negative value;

Suspect: TB – 12/1 00:15 3:00 5:00 12:15 20:00 23:15 23:45 12/2 00:45-1:15 2:00-15 3:45-4:00 12/2 13:00 18:30 12/3 00:00 3:00 4:00 21:00 21:30 22:00 22:45 12/4 00:15 00:45 5:30 10:30 18:30 21:00 12/5 10:00 14:45 19:45 20:30 12/6 1:15-30 5:15 5:45 12:00 14:30 16:15 18:45 20:00-15 12/7 00:30 9:00 9:30 11:30 15:45 19:00 20:00 21:00 12/8 00:00 3:30 9:15 18:45 19:30 20:00 21:15 12/9 00:45 4:15 12/9 11:45 18:00 21:00 22:00 23:15 12/10 00:15-30 1:15 1:45-2:00 4:15-30 9:45 18:15 18:45 19:30 20:15 20:45 22:15-30 23:00-15 23:45 12/11 00:45 3:15 4:30 5:00 5:45 15:45 20:15 12/12 00:00 1:45-2:00 2:45 18:30 20:30 21:30 22:45 12/13 4:15 11:30 20:15 12/14 6:15 9:00 18:45 20:45 22:00 12/15 16:00 20:15 21:15 23:00 12/16 4:45 6:15 10:15 16:00 16:45-17:00 18:15 18:45-19:00 19:30-45 20:45-21:00 12/17 15:45 23:15 12/18 1:00 1:30 3:15 6:15 13:15 13:45 14:15 17:30 23:00 12/19 00:30 6:00 15:45 19:00 19:30 21:45 12/20 2:15 8:45-9:00 12:45 18:00 19:15 12/21 1:15 2:00-15 2:45 14:45 12/22 4:30 5:30 16:00 12/23 1:00 3:00 9:00 10:00-15 11:45 12/23 22:45 12/24 2:00-15 5:45 16:45 18:15 18:45 23:30 12/25 2:00 3:15 4:00 18:45 21:30 22:30 12/26 3:00 4:45 10:15 18:15 20:15 12/27 10:15 12:15 13:30 12/28 4:45 13:30 14:30 12/29 18:15 22:00-15 12/30 8:00 12/30 10:30 14:00  anomalous single point spike; TAL – 11/25 22:15-12/11 1:15 passed Cal ICV low range CCV but failed high range CCV,

Notes: TB – mount is placed in seagrass bed bellow seagrass canopy height, large lizard fish commonly lays eggs inside sonde mount possibly leading to the large amount of spikes observed.

BBMRDW

Missing:

Reject:

Suspect: TB – 12/10 19:00 12/15 14:30 12/17 14:00 12/29 12:30 anomalous single point spike; TAL – 12/23 10:30-1/8 10:30 passed Cal ICV low range CCV but failed high range CCV by small margin;

Notes:

BBMRSC

Missing:

Reject:

Suspect:

Notes:

BBCWA4

Missing:

Reject:

Suspect:

Notes:

BBMRRB

Missing:

Reject:

Suspect: TB – 12/8 2:45 12/31 7:45 anomalous single point or group spike; TAL – 11/13 13:00-12/10 16:15 passed Cal ICV low range CCV but failed high range CCV, 12/10 16:30-1/6 10:45 passed Cal ICV high range CCV but failed low range CCV,

Notes:

15) Acknowledgement:

The data included with this document were collected by the staff of the Florida Department of Environmental Protection at Biscayne Bay Aquatic Preserves. Any products derived from these data should clearly acknowledge this source (please use the attached logos). This recognition is important for ensuring that this long-term monitoring program continues to receive the necessary political and financial support.