St. Martins Marsh and Big Bend Seagrasses Aquatic Preserves (SMMAP & BBSAP)

**Water Quality Metadata Report**

January - December 2019  
Latest Update: July 19, 2021

Note: This is a provisional metadata document; it has not been authenticated as of its download date. Contents of this document are subject to change throughout the QAQC process and it should not be considered a final record of data documentation until that process is complete. Contact the Aquatic Preserve office [Timothy.W.Jones@dep.state.fl.us](mailto:Timothy.W.Jones@dep.state.fl.us) with any additional questions.

**I. Data Set and Research Descriptors**

**1) Principal investigator & contact persons-**

Principal Investigator:

Trisha Green, B.S., Environmental Specialist I

3266 N. Sailboat Ave

Crystal River, FL 34428

(352) 228-6033

Trisha.Green@dep.state.fl.us

Other Contact Persons:

Jamie Letendre, Environmental Specialist II

(352) 228-6032

Jamie.Letendre@dep.state.fl.us

Timothy Jones, Environmental Specialist III, Manager

(352) 228-6031

Timothy.W.Jones@dep.state.fl.us

**2) Entry verification-**

Deployment data are uploaded from the YSI data sonde to a Personal Computer (IBM compatible). Files are exported from EcoWatch in a comma-delimited format (.CDF), EcoWatch Lite in a comma separated file (.CSV) or KOR Software in an Excel File (.XLS). Copies of these files and calibration and field logs and saved in the APWQ Data folder on the common drive, and then the Data Manager is notified by email of the new files. The Data Mangaer removes pre- and post- deployment readings and save the files in the limited.csv folder. Those files are uploaded to the NERRS Centralized Data Management Office (CDMO) Non-SWMP Data Upload Service where data undergo automated primary QAQC. Where deployment overlap occurs between files, the data produced by the newly calibrated sonde is generally accepted as being the most accurate. For more information on QAQC flags and codes, see Sections 11 and 12. During primary QAQC, data are flagged if they are missing or out of sensor range. Primary QAQC files are merged into monthly files in the monthly.csv folder. The files are then returned to the Aquatic Preserve’s Principal Investigator for secondary QAQC where they are analyzed for malfunctions, suspect and/or anomalous data, and notes are made. Files are merged into monthly, quarterly, and annual files and saved on the common drive. Quarterly and annual metadata reports are also saved in the same folder. The Data Manager applies QAQC flags and codes to the quarterly data files using the Aquatic Preserve’s Principal Investigator’s notes and metadata reports. Upload after secondary QAQC results in ingestion into the Aquatic Preserve database as provisional data. The Data Manager reviews the annual files, metadata reports, calibration logs, and post-deployment checks; creates a document with suggested edits, emails it to the Aquatic Preserve’s Principal Investigator, and saves it in the tertiary QAQC folder on the common drive. A meeting is conducted between the Data Manager and Principal Investigator to make corrections and finalize the data and metadata. After tertiary QAQC, the data and metadata are assimilated into the Aquatic Preserve database as authenticated data.

**3) Research objectives-**

The objective of this effort is to establish baseline data by quantifying the spatial/temporal variability and trends, both seasonally and as a function of tidal force, of selected abiotic parameters within the Aquatic Preserves; to record changes in water quality due to major storm events such as hurricanes; and to use this water quality data to complement the annual seagrass monitoring conducted by the Aquatic Preserves.

**4) Research methods-**

Historically, YSI 600 OMS data sondes had been continuously operated (data collection interval: 30 minutes) at the Crystal River site since January 2004, at the Bennett Creek, Kings Bay, and Homosassa River monitoring stations since February 2004, and at the Withlacoochee monitoring station since March 2004. YSI 6600 EDS data sondes were operated at Cat Island and Lone Cabbage Key beginning in March 2004, and at Seahorse Key and Gomez Rocks beginning in April 2004. These models incorporate a specially designed wiper apparatus attached to the turbidity probe that reduces the oxygen and pH sensor fouling and thereby improves the quality of data collected. At each site, the sonde is contained within a 10 cm (inside diameter) housing pipe mounted vertically on a piling. To facilitate water flow across the sensors, several 2 cm diameter holes were drilled into the submerged portion of the pipe. Hole density is greatest near the base where the sonde sensors are located. In early 2005, the Cat Island and Lone Cabbage Key sites were removed. During most of 2005 and early 2006, data was sporadically collected at all sites due to lack of staff.

As of March 2006, all YSI 600 OMS sites were operational. In July 2006, the Gomez Rocks site was removed, and in August 2006, the Seahorse Key station became operational. In October 2006, the YSI 600 OMS was replaced with an YSI 6600 EDS sonde at the Kings Bay station. In March 2007, a 6600 EDS station was established in Dekle Beach. In early 2009, the data collection interval was changed to 15 minutes at all locations. In March 2009, an additional 6600 EDS station was installed at the mouth of the Suwannee River. In March 2010, the Kings Bay station was relocated due to the replacement of the piling the sonde was previously located at. In February 2012, all four YSI 6600 EDS sondes were upgraded from rapid pulse dissolved oxygen probes to ROX optical dissolved oxygen probes. In May 2015, the Crystal River site was broken down due to piling replacement, and the Seahorse Key site was deconstructed. The Homosassa site was upgraded from a YSI 600 OMS to a YSI 6600 EDS in August of 2015. The Kings Bay location was downgraded from a YSI 6600 EDS to a YSI 600 OMS outfitted with a turbidity probe. Due to insufficient staffing, data from 2015 to 2017 are intermittent, and all sites were decommissioned in 2017.

In July 2018, a new station was installed in Chassahowitzka. A 6600 EDS is being used with a 15-minute data collection interval. Parameters being recorded include time, date, temperature (°C), specific conductivity (mS/cm), salinity (ppt), dissolved oxygen (% and mg/L), depth (m), pH, and turbidity (NTU). The data sonde tube is attached to a piling with hose clamps, and water flow through the tube is facilitated with a series of drilled out 2-in and 1-in holes in the submerged portion of the tube.

At the end of October 2019, a new station was installed in Steinhatchee. An EXO2 is being used with a 15-minute data collection interval. Parameters being recorded include time, date, temperature (°C), specific conductivity (mS/cm), salinity (psu), dissolved oxygen (% and mg/L), depth (m), pH, turbidity (FNU), and chlorophyll (RFU and µg/L). The data sonde tube is attached to a channel marker with hose clamps, and water flow through the tube is facilitated with a series of drilled out 2-in and 1-in holes in the submerged portion of the tube. This site was created to fill a data gap that exists in the Big Bend and was funded by an EPA grant awarded to BBSAP in 2019.

Sonde exchanges at the 6600 EDS sites are made at approximately two-week intervals, and sonde exchanges at the EXO sites are made at approximately 30-day intervals. At the end of a sampling period, sondes are returned to the laboratory where post-deployment readings and, if necessary, reconditioning take place in accordance with the methods outlined in the YSI Operating and Service Manual. The turbidity wiper brush is removed and replaced with a clean wiper to avoid contamination of standards during post-deployment procedures. After a superficial rinse of the sonde in tap water, post deployment readings are recorded for pH (RICCA 7.00 buffer solution) and specific conductivity (RICCA 50.00 mS/cm standard). A post-deployment turbidity reading in 0.0 NTU standard (DI water) is recorded after a more thorough rinse of the turbidity sensor. The results of these post-deployment readings are used to evaluate the validity of data (See Table 2).

**5) Site locations and character-**

The St. Martins Marsh Aquatic Preserve was established on October 21, 1969. The St. Martins Marsh Aquatic Preserve covers open water areas from the Crystal River to the Homosassa River in coastal Citrus County. It is composed of approximately 28,400 acres of open water, several inlet bays, tidal rivers and creeks, salt marsh, and adjoins upland hammock islands. Nutrient exchange between the marshes and the Gulf of Mexico make the salt marsh a significant area of primary production and a nursery ground for commercial and recreational fish species. St. Martins Marsh Aquatic Preserve’s freshwater tributaries includes two, first-magnitude, spring-fed rivers: the Homosassa River to the south and the Crystal River to the north. Spring discharge does not fluctuate dramatically from season to season allowing a constant flow of freshwater into St. Martins Marsh’s productive and well-balanced estuary. The area’s vast coastal salt marshes, mud flats, oyster bars, mangrove islands, and seagrass beds are the southern terminus for migratory waterfowl of the Atlantic and Mississippi flyways. St. Martins Marsh provides stop-over and wintering areas for many migratory species. The Springs Coast is characterized by unique limestone outcroppings and exposed karstic features. Habitats associated with these areas are seagrass meadows and hardbottom. Hardbottom habitat is defined as an area of hard substrate, natural or artificial, where macroalgae, sponges, and corals can grow and attach using specialized holdfasts. See Table 1 for a description of the Chassahowitzka datalogger site in St. Martins Marsh Aquatic Preserve.

The Big Bend Seagrasses Aquatic Preserve was established in 1985. Its boundaries extend from the Withlacoochee River north to the St. Marks River and out nine nautical miles. The Preserve boundary encompasses all tidal lands, islands, seagrass beds, shallow banks, and submerged bottoms from the mean high-water line. Landward, it includes all-natural waterways tidally connected to the preserve to the extent of state jurisdiction. Spanning over 945,000 acres, the Big Bend Seagrasses Aquatic Preserve is the largest aquatic preserve and one of the most pristine places in Florida. The Big Bend Seagrasses Aquatic Preserve consists mainly of a large, remote, and undeveloped expanse of submerged seagrasses and nearshore marshlands located along approximately 180 miles of the northeast coast of the Gulf of Mexico where the Florida peninsula joins the panhandle. Numerous estuaries, which nurture a diverse flora and fauna, are formed at the confluence of the many rivers and streams that flow into the Preserve. Open waters and submerged bay bottoms of these estuaries provide habitat to a wide variety of sea and shore birds. This region supports a very important commercial shellfish industry including Cedar Key clams, scallops, oysters, pink shrimp, and blue crab. This area of Florida is also a popular destination for the recreational scallop season. The Suwannee River region supports Essential Fish Habitat (EFH) and the most viable population of the threatened Gulf sturgeon. Big Bend’s vast seagrass beds with mud and sand substrates are important marine habitats to this species. See Table 1 below for a description of the Steinhatchee datalogger site in Big Bend Seagrasses Aquatic Preserve.

**Table 1: Station Descriptions**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **Station Code** | **Site Name** | **Location** | **Active Dates** | **Reason Decommissioned** | **Notes** |
| BBSCH | Chassahowitzka | N 28.77514  W 82.71631 | 07/2018- Present | N/A | Sand/mud bottom, near mouth of Homosassa River |
| BBSST | Steinhatchee | N 29.6625111  W 83.4289231 | 10/2019- Present | N/A | Mud bottom near the mouth of the Steinhatchee River |

**6) Data collection period-**

Individual sonde deployment and retrieval dates and times for 2019 data are as follows:

BEGAN ENDED

Chassahowitzka Site

01/14/2019, 09:00 01/29/2019, 10:15

01/29/2019, 10:30 02/13/2019, 10:15

02/13/2019, 10:30 02/27/2019, 10:15

02/27/2019, 10:30 03/15/2019, 09:00

03/15/2019, 09:15 04/02/2019, 10:00

04/02/2019, 10:15 04/17/2019, 09:15

04/17/2019, 09:30 05/02/2019, 08:45

05/02/2019, 09:00 05/16/2019, 08:15

05/16/2019, 08:30 06/03/2019, 08:30

06/03/2019, 08:45 06/21/2019, 08:15

06/21/2019, 08:30 07/08/2019, 07:15

07/08/2019, 07:30 07/23/2019, 10:15

07/23/2019, 10:30 08/06/2019, 08:30

08/06/2019, 08:45 08/21/2019, 07:45

08/21/2019, 08:00 09/03/2019, 08:15

09/03/2019, 08:30 09/16/2019, 08:30

09/16/2019, 08:45 10/01/2019, 08:30

10/01/2019, 08:45 10/17/2019, 09:15

10/17/2019, 09:30 10/31/2019, 07:30

10/31/2019, 07:45 11/18/2019, 10:00

11/18/2019, 10:15 12/03/2019, 09:30

12/03/2019, 09:45 12/20/2019, 11:15

12/20/2019, 11:30 01/03/2019, 09:00

Steinhatchee Site

10/28/2019, 10:30 11/27/2019, 09:45\*

11/27/2019, 10:00 12/05/2019, 11:30

12/05/2019, 11:45 01/02/2019, 09:15

\* indicates short term loss of data due to battery failure, out of water for maintenance, weather related causes, and/or other internal problems that occurred during deployment.

\*\* indicates long term loss of data due to sonde removal from field for long term repairs.

**7) Distribution-**

The Principle Investigator (PI) retains the right to be fully credited for having collected and processed the data. Following academic courtesy standards, the PI and Aquatic Preserve (AP) site where the data were collected will be contacted and fully acknowledged in any subsequent publications in which any part of the data are used. The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement. The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons. Water quality data and metadata can be obtained from the PI (please see Principal Investigators and Contact Persons) and online at the Aquatic Preserves data portal home page [www.floridaapdata.org](http://www.floridaapdata.org). Data are available in comma delimited format.

**8) Associated researchers and projects-**

The SMMAP and BBSAP have formed partnerships with other agencies and organizations actively involved in resource protection in the Preserves’ watershed. Cooperating managers of lands within the AP’s include the: 1) National Park Service (NPS), 2) Suwannee River Water Management District (SRWMD), 3) Southwest Florida Water Management District (SWFWMD), 4) Department of Environmental Protection (DEP) Division of Recreation and Parks/Florida Park Service (FPS), 5) United States Fish and Wildlife Service (USFWS), 6) DEP Aquatic Preserve Program, 7) Florida Fish and Wildlife Conservation Commission (FWC), 8) Florida Forest Service (FFS), and 9) Citrus, Taylor, Jefferson, Dixie, Levy, and Wakulla Counties.

Other water quality research and monitoring initiatives within the Aquatic Preserves include: 1) Nutrient sampling in the Big Bend region in conjunction with DEP’s Division of Environmental Assessment and Restoration (DEAR) and 2) Project COAST in conjunction with the laboratory of Dr. Thomas Frazer at the University of Florida (This work includes water quality, light penetration (PAR), and nutrient analyses.).

**II. Physical Structure Descriptors**

**9) Sensor specifications-**

6600 EDS data sondes were deployed at the Chassahowitzka site for all of 2019, and EXOs were deployed at Steinhatchee.

**YSI 6600 EDS Datasonde:**

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Thermistor

Model#: 6560

Range: -5 to 50 C

Accuracy: +/- 0.15

Resolution: 0.01 C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: 4-electrode cell with autoranging

Model#: 6560

Range: 0 to 100 mS/cm

Accuracy: +/- 0.5% of reading + 0.001 mS/cm

Resolution: 0.001 mS/cm to 0.1 mS/cm (range dependant)

Parameter: Salinity

Units: parts per thousand (ppt)

Sensor Type: Calculated from conductivity and temperature

Range: 0 to 70 ppt

Accuracy: +/- 1.0% of reading or 0.1 ppt, whichever is greater

Resolution: 0.01 ppt

Parameter: Dissolved Oxygen % saturation

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 6150 ROX

Range: 0 to 500% air saturation

Accuracy: 0-200% air saturation: +/- 1% of the reading or 1% air saturation, whichever is greater 200-500% air saturation: +/- 15% or reading

Resolution: 0.1% air saturation

Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature, and salinity)

Units: milligrams/Liter (mg/L)

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 6150 ROX

Range: 0 to 50 mg/L

Accuracy: 0-20 mg/L: +/-0.1 mg/l or 1% of the reading, whichever is greater

20 to 50 mg/L: +/- 15% of the reading

Resolution: 0.01 mg/L

Parameter: Non-vented Level - Shallow (Depth)

Units: feet or meters (ft or m)

Sensor Type: Stainless steel strain gauge

Range: 0 to 30 ft (9.1 m)

Accuracy: +/- 0.06 ft (0.018 m)

Resolution: 0.001 ft (0.001 m)

Parameter: pH – bulb probe or EDS flat glass probe

Units: pH units

Sensor Type: Glass combination electrode

Model#: 6561 or 6561FG

Range: 0 to 14 units

Accuracy: +/- 0.2 units

Resolution: 0.01 units

Parameter: Turbidity

Units: nephelometric turbidity units (NTU)

Sensor Type: Optical, 90-degree scatter, with mechanical cleaning

Model#: 6136

Range: 0 to 1000 NTU

Accuracy: +/- 2% of reading or 0.3 NTU (whichever is greater)

Resolution: 0.1 NTU

**YSI EXO Datasonde:**

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Wiped probe; Thermistor

Model#: 599827

Range: -5 to 50 C

Accuracy: ±0.2 C

Resolution: 0.001 C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: Wiped probe; 4-electrode cell with autoranging

Model#: 599827

Range: 0 to 100 mS/cm

Accuracy: ±1% of the reading or 0.002 mS/cm, whichever is greater

Resolution: 0.0001 to 0.01 mS/cm (range dependent)

Parameter: Salinity

Units: practical salinity units (psu)/parts per thousand (ppt)

Model#: 599827

Sensor Type: Wiped probe; Calculated from conductivity and temperature

Range: 0 to 70 ppt

Accuracy: ±2% of the reading or 0.2 ppt, whichever is greater

Resolution: 0.01 psu

Parameter: Dissolved Oxygen % saturation

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100-01

Range: 0 to 500% air saturation

Accuracy: 0-200% air saturation: +/- 1% of the reading or 1% air saturation, whichever is greater 200-500% air saturation: +/- 5% or reading

Resolution: 0.1% air saturation

Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature, and salinity)

Units: milligrams/Liter (mg/L)

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100-01

Range: 0 to 50 mg/L

Accuracy: 0-20 mg/L: +/-0.1 mg/l or 1% of the reading, whichever is greater

20 to 50 mg/L: +/- 5% of the reading

Resolution: 0.01 mg/L

Parameter: Non-vented Level - Shallow (Depth)

Units: feet or meters (ft or m)

Sensor Type: Stainless steel strain gauge

Range: 0 to 33 ft (10 m)

Accuracy: +/- 0.013 ft (0.004 m)

Resolution: 0.001 ft (0.001 m)

Parameter: pH

Units: pH units

Sensor Type: Glass combination electrode

Model#: 599701(guarded) or 599702(wiped)

Range: 0 to 14 units

Accuracy: +/- 0.1 units within +/- 10° of calibration temperature, +/- 0.2 units for entire temperature range

Resolution: 0.01 units

Parameter: Turbidity

Units: formazin nephelometric units (FNU)

Sensor Type: Optical, 90 degree scatter

Model#: 599101-01

Range: 0 to 4000 FNU

Accuracy: 0 to 999 FNU: 0.3 FNU or +/-2% of reading (whichever is greater); 1000 to 4000 FNU +/-5% of reading

Resolution: 0 to 999 FNU: 0.01 FNU, 1000 to 4000 FNU: 0.1 FNU

Parameter: Chlorophyll

Units: micrograms/Liter

Sensor Type: Optical probe

Model#: 599102-01

Range: 0 to 400 ug/Liter

Accuracy: Dependent on methodology

Resolution: 0.1 ug/L chl a, 0.1% FS

**Depth Qualifier:**

YSI data sondes can be equipped with either vented or non-vented depth/level sensors.  Readings for both vented and non-vented sensors are automatically compensated for water density change due to variations in temperature and salinity; but for all non-vented depth measurements, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.02 cm for every 1 millibar change in atmospheric pressure and is eliminated for vented sensors because they are vented to the atmosphere throughout the deployment time interval.

Standard calibration protocol calls for all non-vented depth sensors to read 0 meters at a (local) barometric pressure of 1013.25 mb (760 mm/hg).  To achieve this, each site calibrates their depth sensor with a depth offset number, which is calculated using the actual atmospheric pressure at the time of calibration and the equation provided in the Aquatic Preserve calibration sheet or digital calibration log.  This offset procedure standardizes each depth calibration. If accurate atmospheric pressure data are available, non-vented sensor depth measurements can be corrected. The Principal Investigator should be contacted in order to obtain information regarding atmospheric pressure data availability.

**Salinity Units Qualifier:**

The 6600 series sondes report salinity in parts per thousand (ppt) units, and the EXO sondes report practical salinity units (psu). These units are essentially the same and for Aquatic Preserve purposes are understood to be equivalent; however, psu is considered the more appropriate designation. Moving forward, the Aquatic Preserve program will assign psu salinity units for all data regardless of sonde type.

**Turbidity Qualifier:**

The 6600 series sondes report turbidity in nephelometric turbidity units (NTU), the EXO sondes use formazin nephelometric units (FNU). These units are essentially the same but indicate a difference in sensor methodology, for Aquatic Preserve purposes they will be considered equivalent. Moving forward, the Aquatic Preserve program will use FNU/NTU as the designated units for all turbidity data regardless of sonde type. If turbidity units and sensor methodology are of concern, please see the Sensor Specifications portion of the metadata.

**Chlorophyll Fluorescence Disclaimer:**

YSI chlorophyll sensors (6025 or 599102-01) are designed to serve as a proxy for chlorophyll concentrations in the field for monitoring applications and complement traditional lab extraction methods; therefore, there are accuracy limitations associated with the data that are detailed in the YSI manual including interference from other fluorescent species, differences in calibration method, and effects of cell structure, particle size, organism type, temperature, and light on sensor measurements.

**10) Coded variable definitions**

Sampling station: Sampling site code: Station code:

Chassahowitzka CH BBSCH

Steinhatchee ST BBSST

**11) QAQC flag definitions-**

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F\_). During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is missing and above or below sensor range. All remaining data are then flagged 0, passing initial QAQC checks. During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5 Outside High Sensor Range

-4 Outside Low Sensor Range

-3 Data Rejected due to QAQC

-2 Missing Data

-1 Optional SWMP Supported Parameter

0 Data Passed Initial QAQC Checks

1 Suspect Data

2 *Open - reserved for later flag*

3 Calculated data: non-vented depth/level sensor correction for changes in barometric pressure

4 Historical Data: Pre-Auto QAQC

5 Corrected Data

**12) QAQC code definitions**

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column. There are three (3) different code categories, general, sensor, and comment. General errors document general problems with the deployment or YSI datasonde, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data. Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an \* below) can be applied to the entire record in the F\_Record column.

General Errors

GIC No instrument deployed due to ice

GIM Instrument malfunction

GIT Instrument recording error; recovered telemetry data

GMC No instrument deployed due to maintenance/calibration

GNF Deployment tube clogged / no flow

GOW Out of water event

GPF Power failure / low battery

GQR Data rejected due to QA/QC checks

GSM See metadata

Corrected Depth/Level Data Codes

GCC Calculated with data that were corrected during QA/QC

GCM Calculated value could not be determined due to missing data

GCR Calculated value could not be determined due to rejected data

GCS Calculated value suspect due to questionable data

GCU Calculated value could not be determined due to unavailable data

Sensor Errors

SBO Blocked optic

SCF Conductivity sensor failure

SCS Chlorophyll spike

SDF Depth port frozen

SDG Suspect due to sensor diagnostics

SDO DO suspect

SDP DO membrane puncture

SIC Incorrect calibration / contaminated standard

SNV Negative value

SOW Sensor out of water

SPC Post calibration out of range

SQR Data rejected due to QAQC checks

SSD Sensor drift

SSM Sensor malfunction

SSR Sensor removed / not deployed

STF Catastrophic temperature sensor failure

STS Turbidity spike

SWM Wiper malfunction / loss

Comments

CAB\* Algal bloom

CAF Acceptable calibration/accuracy error of sensor

CAP Depth sensor in water, affected by atmospheric pressure

CBF Biofouling

CCU Cause unknown

CDA\* DO hypoxia (<3 mg/L)

CDB\* Disturbed bottom

CDF Data appear to fit conditions

CFK\* Fish kill

CIP\* Surface ice present at sample station

CLT\* Low tide

CMC\* In field maintenance/cleaning

CMD\* Mud in probe guard

CND New deployment begins

CRE\* Significant rain event

CSM\* See metadata

CTS Turbidity spike

CVT\* Possible vandalism/tampering

CWD\* Data collected at wrong depth

CWE\* Significant weather event

**13) Post deployment information**

End of deployment post-calibration readings in standard solutions are taken prior to probe cleaning.

**Table 2. Post-deployment readings of 6600 EDS sondes deployed at the Chassahowitzka site** **during 2019.**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | **pH** | **SpCond (mS/cm)** | **DO %** | **Turbidity (NTU)** | **Depth (m)** |
| **Date/Std.** | **7.00** | **50.00** | **100.0** | **0.0** | **n/a** |
| 01/14/2019 | 6.94 | 47.76 | 100.3, 100.4 | -0.1 | 0.138 |
| 01/29/2019 | 7.00 | 26.40 | 47.1, 43.7 | -0.4 | 0.017 |
| 02/13/2019 | 7.01 | 50.38 | 99.7, 99.4 | -0.1 | 0.069 |
| 02/27/2019 | 7.04 | 50.19 | 100.5, 99.8 | 0.0 | 0.078 |
| 03/15/2019 | 7.30 | 50.87 | 100.4, 102.3 | -0.3 | 0.067 |
| 04/02/2019 | 7.17 | 50.60 | 100.5, 99.5 | 0.0 | -0.013 |
| 04/17/2019 | 7.27 | 50.21 | 101.0, 100.6 | 0.1 | 0.063 |
| 05/02/2019 | 7.08 | 50.67 | 99.7, 103.2 | 0.2 | 0.007 |
| 05/16/2019 | 6.92 | 49.67 | 99.0, 98.9 | 0.1 | 0.026 |
| 06/03/2019 | 6.96 | 50.83 | 101.5, 102.3 | 0.1 | 0.053 |
| 06/21/2019 | 6.96 | 50.48 | 99.4, 99.4 | 0.2 | 0.000 |
| 07/08/2019 | 6.99 | 50.61 | 99.9, 99.0 | 0.0 | 0.001 |
| 07/23/2019 | 7.28 | 51.17 | 99.9, 99.8 | -5.1 | 0.014 |
| 08/06/2019 | 7.15 | 50.67 | 100.8, 100.8 | 0.1 | 0.041 |
| 08/21/2019 | 7.52 | 50.38 | 99.0, 98.9 | -0.2 | -0.032 |
| 09/03/2019 | 6.85 | 50.40 | 100.1, 100.3 | 0.0 | 0.036 |
| 09/16/2019 | 7.00 | 50.40 | 99.0, 99.0 | N/A | 0.017 |
| 10/01/2019 | 7.16 | 50.19 | 99.5, 99.5 | 0.1 | -0.033 |
| 10/17/2019 | 7.05 | 50.11 | 100.3, 100.3 | -0.2 | 0.069 |
| 10/31/2019 | 6.97 | 32.08 | 67, 69.2 | -0.1 | -0.045 |
| 11/18/2019 | 7.77 | 49.75 | 101.7, 101.4 | 0.1 | 0.047 |
| 12/03/2019 | 6.83 | 50.05 | 101.5, 101.8 | -0.3 | 0.055 |
| 12/20/2019 | 7.20 | 50.99 | 99.7, 100.3 | -0.3 | 0.049 |

**Table 3. Post-deployment readings of EXO2 sondes deployed at the Steinhatchee site** **during 2019.**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | **pH** | **SpCond (mS/cm)** | **DO %** | **Turbidity (FNU)** | **Depth (m)** |
| **Date/Std.** | **7.00** | **50.00** | **100.0** | **0.0** | **n/a** |
| 10/28/2019 | 7.16 | 49.632 | 100.6, 100.6 | 1.01 | 0.041 |
| **\***11/27/2019 | 7.01 | 49.766 | 99.9, 99.8 | 1.09 | 0.035 |
| 12/05/2019 | 7.02 | 49.675 | 99.8, 100.4 | 0.13 | 0.015 |

**\***EXO3 was briefly deployed due to second EXO2 at YSI for repair.

**14) Other remarks/notes:**

1. Calibration of dissolved oxygen was performed on the day of deployment. Two dissolved oxygen values are recorded during the post-deployment evaluation process.
2. This style of metadata was formerly used by National Estuarine Research Reserve program; more specifically, this report was modeled after metadata reports created at the Guana Tolomato Matanzas National Estuarine Research Reserve.
3. Copies of calibration/deployment logs can be obtained through the Principal Investigator.
4. Accreditation must be given to Florida Department of Environmental Protection’s Office of Resilience and Coastal Protection staff of the St. Martins Marsh and Big Bend Seagrasses Aquatic Preserves for all data used.
5. Between August 13 and August 18, 2019, it was estimated over 7 inches of rain fell in Citrus County with some areas reaching 9-11 inches of precipitation.
6. A turbidity probe was not available to install on one of the 6600’s at the Chassahowitzka site for the 09/16/2019 deployment. An old 2014 turbidity probe was installed for the 10/17/2019 deployment.
7. An EXO3 was deployed for about eight days at the Steinhatchee site to relieve the EXO2 from rotation while it was post-calibrated, cleaned, and re-calibrated for deployment again. The second EXO2 was at YSI for repair, so it was not available for swapping. No chlorophyll data was collected during this time; the EXO3 doesn’t have enough sensor ports to include a chlorophyll sensor.

**Missing data**

**Station BBSCH:**

**September 1-30, 2019**

**CH**

1. Missing turbidity data 09/16/2019 08:45- 09/30/2019 23:45; deployed without a turbidity probe. A turbidity probe was not available to replace previously failed turbidity probe.

**October 1-31, 2019**

**CH**

1. Missing turbidity data 10/01/2019 00:00-09:45; deployed without a turbidity probe. A turbidity probe was not available to replace previously failed turbidity probe.

**Station BBSST:**

**November 1-31, 2019**

**ST**

1. Missing chlorophyll data 11/27/2019- 11/30/2019 23:45; EXO3 had to be deployed while new EXO2 went in for service. EXO3 doesn’t have enough ports for a chlorophyll sensor.

**December 1-31, 2019**

**ST**

1. Missing chlorophyll data 12/01/2019 00:00- 12/05/2019 11:30; EXO3 had to be deployed while new EXO2 went in for service. EXO3 doesn’t have enough ports for a chlorophyll sensor.

**Anomalous/suspect/rejected data**

**Note #1:** Slight shifts in data are sometimes correlated with sonde exchanges. These shifts are most noticeable in pH, specific conductivity, salinity, DO% and DO conc, and may be related to sensor drift (e.g., due to fouling) and/or calibration/performance differences between sondes.

**Note #2:** Turbidity “outliers” (i.e., values that are negative or greater than 1000 NTU for 6600 series sondes and 4000 NTU for EXO series sondes) were not deleted from the monthly records. Readings greater than 1000 NTU for 6600 series sondes and 4000 NTU for EXO series sondes are considered out of range and are rejected. They have been left in the database to provide users with a complete dataset and to allow true visual representation of the data in graphs. Negative turbidity values occur throughout the year at all sites. Some of these negative values are within the accuracy range of the sensor (+/- 2.0 %) and, therefore, were not removed from the dataset. They were marked suspect with the CAF code.

**Note #3**: Turbidity data is subject to single and clusters of spikes that occur in the beginning and middle of deployments. Turbidity values that fall between 500 and 1000 are not specifically indicated as suspect data, but possibly could be interpreted as suspect. Turbidity spikes may be associated with wiper malfunction but mostly the reason is unknown. Data users should exercise caution when interpreting turbidity data that fall within this range.

**Note #4:** Time series profiles of the dissolved oxygen data at all monitoring stations sometimes exhibits brief “spikes” of reduced DO concentrations. These events appear to be coupled with the occurrence of slack tide conditions as well as the level of fouling associated with the sonde.

**Note #5:** All times in data files at all sites had to be adjusted post-deployment; more specifically, times were altered such that the readings occurred on the hour and half hour. It has been determined that this clock error is a software issue and has been resolved.

**Station BBSCH:**

**January 1-31, 2019**

**CH**

1. Rejected all data 01/30/2019 20:30- 01/31/2019 23:45 due to catastrophic temperature probe failure.

**February 1-28, 2019**

**CH**

1. Rejected all data 02/01/2019 00:00- 02/13/2019 10:30 due to catastrophic temperature probe failure.
2. Suspect all data 02/27/2019 10:30- 03/15/2019 09:00 due to unmaintained temperature probe that experienced failure (Probe #14E 100327). Data appear to fit conditions.

**March 1-31, 2019**

**CH**

1. Suspect negative turbidity data 03/23/2019 16:45; may be contributed to calibration offset.

**April 1-30, 2019**

**CH**

1. Suspect turbidity data 04/24/2019 14:30; value outside of high sensor range during high turbidity spike.
2. Suspect all data 04/02/2019 10:15- 04/17/2019 09:15 due to unmaintained temperature probe that experienced failure (Probe #14E 100327). Data appear to fit conditions.

**May 1-31, 2019**

**CH**

1. Suspect all data 05/02/2019 09:00- 05/ due to unmaintained temperature probe that experienced failure (Probe #14E 100327). Data appear to fit conditions.

**June 1-30, 2019**

**CH**

1. Suspect turbidity data 06/21/2019 9:00- 06/27/2019 17:45, 06/28/2019 02:30-15:00, 15:30-17:30, 18:45, 06/28/2019 19:30- 06/29/2019 00:00-15:15, 16:15, 17:00-18:30, 06/29/2019 19:15- 6/30/2019 23:45; data appear to fit conditions but are marked suspect due to intermittent turbidity probe malfunction.
2. Rejected depth data 06/03/2019 08:45- 06/21/2019 08:15; shift in depth data that corresponds to a datasonde exchange; may be due to temperature probe malfunction.
3. Rejected turbidity data 06/21/2019 08:45, 06/27/2019 18:00- 06/28/2019 02:15, 15:15, 17:45-18:30, 19:00-19:15, 06/29/2019 15:30-16:00, 16:30-16:45, 18:45-19:00; outside low sensor range. Unknown reason for probe malfunction. Turbidity calibrated and post-calibrated within acceptable range.
4. All data rejected 06/03/2019 08:45- 06/21/2019 08:15; small jump in temperature data at sonde exchange; same temperature probe (14E 100327) that experienced catastrophic failure 01/02/2019; probe was replaced 12/03/2019.

**July 1-31, 2019**

**CH**

1. Suspect negative turbidity data 07/01/2019 00:00-19:45, 07/01/2019 20:15-07/03/2019 01:30, 02:00, 04:30, 05:00-5:15, 07/03/2019 07:45- 07/05/2019 03:30, 04:15, 06:15-06:30, 11:15-22:45, 07/06/2019 06:15-12:00, 16:00, 17:15-21:30, 22:00-22:45, 07/07/2019 00:00- 07/08/2019 07:15, 07/23/2019 10:30-16:00, 17:45-18:45, 20:00-20:15, 23:45, 07/26/2019 00:30-01:00, 01:45, 02:45, 04:45-05:00, 07:45-08:00, 08:30, 07/31/2019 18:00, 18:30-19:00; intermittent turbidity plunges/probe malfunction; data appear to fit conditions.
2. Rejected depth data 07/23/2019 10:30- 07/31/2019 23:45; large shift in depth data values between the 07/23/2019 10:30 sonde exchange. Calibration values were double-checked and confirmed valid. Depth channel was clean. Sonde 120954’s depth values are suspect whereas sonde 120953’s depth values are more reasonable for the Chassahowitzka site.
3. Rejected turbidity data 07/01/2019 20:00, 07/03/2019 01:45, 02:15-04:45, 05:30-07:30, 07/05/2019 03:45-04:00, 04:30-06:00, 06:45-11:00, 07/05/2019 23:00- 07/06/2019 6:00, 12:15-15:45, 16:15-17:00, 21:45, 23:00-23:45; 07/25/2019 16:15-7:30, 19:00-19:45, 20:30, 21:00-23:30, 07/26/2019 00:00-00:15, 01:15-01:30, 02:00-02:30, 03:00-04:30, 05:15-07:30, 08:15, 07/26/2019 08:45- 07/31/2019 17:45, 18:15, 19:15-23:45; turbidity probe failure.
4. All data rejected 07/08/2019 07:30- 07/23/2019 10:15; small jump in temperature data at sonde exchange

**August 1-31, 2019**

**CH**

1. Suspect negative turbidity data 08/01/2019 15:45, 08/02/2019 17:00, 08/04/2019 15:00; may be contributed to calibration offset or related to the future intermittent turbidity plunges/probe malfunction.
2. Rejected depth data 08/01/2019 00:00- 08/06/2019 08:30; large shift in depth data between the 07/23/2019 10:30 sonde exchange. Calibration values were double-checked and confirmed valid and depth channel was clean; may be onset of depth probe failure (see September note).
3. Rejected depth data 08/21/2019 08:00- 08/31/2019 23:45; large shift in depth data between the 08/21/2019 08:00 sonde exchange. Calibration values were double-checked and confirmed valid. Depth port/channel was clean.
4. Rejected turbidity data 08/01/2019 00:00-15:30; 08/01/2019 16:00- 08/02/2019 16:30, 08/02/2019 17:15- 08/04/2019 14:45, 08/04/2019 15:15- 08/06/2019 08:30; values out of sensor range due to turbidity probe failure.
5. Rejected turbidity data 08/31/2019 00:00; values out of sensor range due to turbidity probe malfunction or failure. Reason uncertain now because this is a new turbidity probe.
6. All data rejected 08/06/2019 08:45- 08/21/2019 08:30; large jump in temperature data at sonde exchange; same temperature probe (14E 100327) that experienced catastrophic failure 01/02/2019; probe was replaced 12/03/2019.

**September 1-30, 2019**

**CH**

1. Rejected depth data 09/01/2019 00:00- 09/16/2019 09:45; large shift in depth data between the 08/21/2019 08:00 sonde exchange. Calibration values were double-checked and confirmed valid. Depth channel was clean. After post-calibration, I called YSI for help. I followed their instructions and cleared the depth calibration. Then, I used three different depth instruments in our boat basin to validate the correct depth and re-calibrate it. The problem does not continue after this.
2. Suspect turbidity data 09/03/2019 04:00, 04:45, 05:30, 06:00; may be contributed to calibration offset or related to the future intermittent turbidity plunges/probe malfunction.
3. Rejected turbidity data 09/01/2019 19:15-19:30, 20:00, 20:30-20:45, 22:15, 09/02/2019 00:00, 00:30-00:45, 01:15, 01:45-02:00, 03:00, 03:30, 09/03/2019 05:45, 06:15-06:30; values of out sensor range due to turbidity probe malfunction.
4. Rejected turbidity data 09/16/2019 08:45- 09/30/2019 23:45; no sensor deployed.
5. All data rejected09/03/2019 08:30- 09/16/2019 08:30; large jump in temperature data at sonde exchange; same temperature probe (14E 100327) that experienced catastrophic failure 01/02/2019; probe was replaced 12/03/2019.

**October 1-31, 2019**

**CH**

1. All data rejected 10/01/2019 10:00- 10/17/2019 09:15, 10/31/2019 07:45-23:45; noticeable temperature shift observed at sonde exchange on 10/31/2019 07:45 and ultimate conductivity/temperature probe failure at the end of the deployment; same temperature probe (14E 100327) that experienced catastrophic failure 01/02/2019; probe was replaced 12/03/2019.
2. Rejected turbidity data 10/01/2019 00:00-09:45; no sensor deployed.
3. Rejected turbidity data 10/22/2019 15:00, 10/23/2019 06:45, 10/24/2019 01:45, 02:15, 10/26/2019 13:30, 16:00, 10/27/2019 13:30, 14:15, 10/28/2019 12:00, 10/29/2019 12:15, 14:15, 15:15-15:30, 16:00, 10/30/2019 17:00; old turbidity probe was installed trying to gather maximum amount of turbidity data rather than leaving a turbidity probe uninstalled each deployment. Reason for negative spikes unknown.
4. Rejected turbidity spike 10/25/2019 03:00 due to ongoing turbidity probe issues; exact cause unknown.
5. Suspect turbidity data 10/29/2019 15:45.

**November 1-30, 2019**

**CH**

1. All data rejected 11/01/2019 00:00- 11/18/2019 10:00; noticeable temperature shift observed at sonde exchange on 10/31/2019 07:45 and ultimate conductivity/temperature probe failure at the end of the deployment; same temperature probe (14E 100327) that experienced catastrophic failure 01/02/2019; probe was replaced 12/03/2019; temperature probe malfunction and post-calibrated out of range.

**December 1-31, 2019**

**CH**

1. Suspect turbidity data 12/05/2019 22:30, 12/06/2019 21:45, 23:00-23:15, 12/07/2019 00:00, 01:00-01:15, 12/09/2019 00:00, 01:00-01:15, 02:15-02:45, 03:30, 12/10/2019 00:00, 02:00-02:30, 12/10/2019 03:00-03:30, 05:00, 12/11/2019 03:30-04:00; may be due to calibration offset.
2. Noticeable pH shift near the 12/03/2019 and 12/20/2019 sonde exchanges. See Note #1.

**Station BBSST:**

**October 1-31, 2019**

**ST**

1. Corrected depth data 10/28/2019 10:30- 10/31/2019 23:45; depth calibrated incorrectly without depth offset.
2. Rejected chlorophyll readings from 10/28/2019 11:30- 10/31/2019 23:45; TAL-PE sensor was deployed with a factory calibration.

**November 1-30, 2019**

**ST**

1. Corrected depth data 11/01/2019 00:00- 11/30/2019 23:45; depth calibrated incorrectly without depth offset.
2. Rejected chlorophyll data from 11/01/2019 00:00- 11/27/2019 09:45; TAL-PE sensor was deployed with a factory calibration.
3. Rejected all data 11/21/2019 10:45-11:00; sonde was removed from water to download data, check battery life, and perform a thorough cleaning on the boat. No sonde was available to swap this one out with, so it continued the same deployment name from 10/28/2019. Two data files exist for this single deployment.

**December 1-31, 2019**

**ST**

1. Corrected depth data 12/01/2019 00:00- 12/31/2019 23:45; depth calibrated incorrectly without depth offset.