Florida Keys National Marine Sanctuary (FKNMS) & Florida Keys Aquatic Preserves (FKAP) Water Quality Metadata

Document Coverage:

Latest Update: 7/31/2026

Note: This is a provisional metadata document; it has not been authenticated as of its download date.  Contents of this document are subject to change throughout the QAQC process, and it should not be considered a final record of data documentation until that process is complete. Contact the Florida Keys National Marine Sanctuary Environmental Manager at Nicholas.Parr@FloridaDEP.gov with any additional questions.

I. Data Set and Research Descriptors

1)  Principal investigator and contact persons –

Dr. Nicholas Parr – Florida Keys National Marine Sanctuary Liaison & Aquatic Preserves Manager

Nicholas.Parr@FloridaDEP.gov

Office: (305) 289-7083

2976 Overseas Highway, 219C, Marathon, FL 33050

Genevieve Schave – Water Quality Protection Program Coordinator

Genevieve.Schave@FloridaDEP.gov

Office: (305) 289-7085

2976 Overseas Highway, 219B, Marathon, FL 33050

2)  Entry verification –

Deployment data are uploaded from the YSI datasonde to a Personal Computer (IBM compatible).  Files are exported from KOR EXO in a comma-separated file (.CSV) and uploaded to the National Estuarine Research Reserve System (NERRS) Centralized Data Management Office (CDMO) Non-System Wide Monitoring Program (SWMP) Data Upload Service where data undergo automated primary QAQC. All pre- and post-deployment data are removed from the file prior to upload.

During primary QAQC, data are flagged if they are missing or out of sensor range.  Edited files are returned to the Office of Resilience and Coastal Protection (RCP) Data Coordinator and/or Aquatic Preserve office for secondary QAQC. Using Microsoft Excel, the files are opened and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes create a metadata worksheet for flagged data, summarized statistics, and graphed data for review.  It allows the user to apply QAQC flags/codes to the data, remove any overlapping deployment data, append files, and export the resulting data files for upload to the Aquatic Preserve (AP) database.

After the secondary QAQC results are complete the file will be incorporated into the AP database as provisional plus data. A final tertiary QAQC will take place by the RCP’s Data Coordinator where the file will then be assimilated into the AP database as authenticated data.  Any deployment overlaps between files will default to the data collected via the newly calibrated sonde. For more information on QAQC flags and codes, see Sections 11 and 12. Please refer to NERRS CDMO Data Management Manual Version 6.6 (http://cdmo.baruch.sc.edu/request-manuals/) for more information regarding the QA/QC procedures.

All data management and QAQC checks are handled by Genevieve Schave and Dr. Nicholas Parr.

3)  Research objectives –

The Florida Department of Environmental Protection, Office of Resilience and Coastal Protection (FDEP ORCP) oversees many waterways and the programs that support them however, the main focus in the Keys is the aquatic preserves, in addition to our presence within the Florida Keys National Marine Sanctuary (FKNMS).

The Florida Keys has two aquatic preserves that are managed via FDEP ORCP. These include:

- Lignumvitae Aquatic Preserve (LKAP) was established in 1969 between Upper and Lower Matecumbe Key which encompasses 7,500 acres of seagrass meadows and deep-water channels. This area is jointly managed with the Lignumvitae Key Botanical State Park.

- Coupon Bight Aquatic Preserve (CBAP) was established in 1969, located on the southern shore of Big Pine Key and encompasses 6,000 acres of seagrass meadows, hard bottom communities, mangrove wetlands, and coral patch reefs.

The Florida Keys National Marine Sanctuary and Protection Act was passed in 1990, bringing about a sanctuary wide management plan that was implemented in July of 1997. The 2,900 square nautical miles of protected waters is jointly managed by the National Oceanic and Atmospheric Administration (NOAA) and the State of Florida. FDEP serves as the lead agency for water quality management and co-chair of the FKNMS Water Quality Protection Program with the U.S. Environmental Protection Program (EPA).

With water quality being one of the public’s top concerns and priorities these datasondes allow us to collect daily data that can be used for long term management planning and emergent action if noticeable decline is observed. Both stations have unique flow patterns that can help us assess both localized and long-term issues. LKAP’s sonde captures the bayside flow exchange from an Everglades dominated outgoing tide to an Atlantic Ocean dominated incoming tide. CBAP’s sonde station is within the Bight and monitors the tidal influence on an area that once hosted robust coral communities. These datasondes collect abiotic parameters every 15 minutes.

4)  Research methods –

Specific conductance, pH, turbidity, depth, chlorophyll and dissolved oxygen are calibrated on the EXO2s. Calibrations are conducted in the FKNMS/FKAP lab by FKNMS/FKAP staff. Specific conductance is a 1-point calibration using aa 50,000 µS/cm standard (Ricca). The ICV is done using 70,000 or 100,000 µS/cm standard (Ricca) and the CCV is done using 50,000 and 70,000 or 100,000 µS/cm standard. pH is a 2-point calibration using pH 7 standard (Ricca) and pH 10 standard (Ricca). The Initial Calibration Verification (ICV) and Continuous Calibration Verification (CCV) are done using one of the 2 standards. Turbidity is a 2-point calibration with 0 FNU/NTU Milli-Q water and 124 FNU standard (YSI). The ICV and CCV are done with both 0 and 124/126 FNU/NTU standards. Chlorophyll is a 2-point calibration with 0 ug/L Milli-Q water and a rhodamine WT dye standard (Kingscote). IVC and CCV are done using one of the 2 standards.

Data is collected in 15-minute intervals. All sondes are calibrated before deployment and a pre-deployment calibration verification (ICV) is done as a QAQC check. Unless otherwise noted in Section 6, within 30 days of deployment datasondes are removed from the water and returned to the lab. After ambient temperature acclimation, post deployment readings are recorded for pH (Fisher 7.00 and 10.00 buffer solutions) and specific conductivity (YSI 50.00 mS/cm standard); a post-deployment turbidity reading in 0.0 NTU standard (DI water) is recorded after a more thorough rinse of the turbidity sensor.  The results of these post-deployment readings are used to evaluate the validity of data. Calibration log documents all calibrations, verifications, standards and lot numbers used in process. Reconditioning and maintenance take place in accordance with the methods outlined in the YSI Operating and Service Manual.

5)  Site location and character –

In October of 2021, the LKAP continuous water quality monitoring site was established. The EXO2 datasonde mount is constructed using a 3-in PVC pipe that is vertically attached to the Lignumvitae Key visitor’s dock. Using wire rope to control the deployments position, the sonde is lowered down the PVC pipe until the sensors and guard are fully exposed to the water flow. The sensors are situated one meter from the bay floor which allows the sensors to always remain submerged while avoiding temporary sedimentation caused by the large amount of fish activity in the area.

| Site name | Lignumvitae Key Aquatic Preserve (FKLK) |
| --- | --- |
| Latitude and longitude | 24.90345, -80.69456 |
| Tidal range | Mean range 0.66 ft; Diurnal range 0.78 ft (NOAA Station 8723824) |
| Salinity range | 32 - 40 ppt |
| Type and amount of freshwater input | Rainwater; Stormwater runoff from Lignumvitae Key |
| Water depth (MLW) | 2.0 m (6.5 ft) |
| Sonde distance from bottom | 1.0 m (3.3 ft) |
| Bottom habitat or type | Muddy, soft sediment with sparse to patchy macro-algae and seagrass (Thalassia testudinum, Syringodium filiforme, and Halodule wrightii all present) |
| Pollutants in area | Stormwater runoff from the surrounding Keys; Marine debris from local boating and brought in by ocean currents; Mainland runoff from the mainland through Florida Bay (Restoration efforts in the Everglades Ecosystem have helped lower nutrient influx throughout surrounding areas) |
| Description of watershed | This site captures the flow through Indian Key channel, which stretches from Florida Bay to the Atlantic Ocean. This channel, on both bayside and oceanside, is surrounded by plush seagrass flats. |

In October of 2023, a sonde site was established in Coupon Bight Aquatic Preserve is mounted to a Seagrass Bank sign south of Little Munson Island. This station is also within 1 miles of the Newfound Harbor Sanctuary Preservation Area (aka Munson Rocks), one of the seven Iconic Reefs of the Florida Keys National Marine Sanctuary. This site is more than 500m from the nearest developed shoreline and therefore is not considered to be heavily influenced by land-based sources of pollution.

| Site name | Coupon Bight Aquatic Preserve (FKCB) |
| --- | --- |
| Latitude and longitude | 24.62045, -81.40563 |
| Tidal range | Mean range 1.33 ft; Diurnal range 1.73 ft (NOAA Station 8724238) |
| Salinity range | 33 – 41 ppt |
| Type and amount of freshwater input | Rainwater |
| Water depth (MLW) | 3.0m (9.9 ft) |
| Sonde distance from bottom | 1.0m (3.3 ft) |
| Bottom habitat or type | On the edge of a channel and seagrass bed, muddy sediment, thick seagrass bed (Thalassia testudinum dominated). Coral patch reefs and mixed hardbottom communities dominate the adjacent areas. |
| Pollutants in area | Marine debris and pollutants from local boating and brought in by ocean currents |
| Description of watershed | This location captures the flow from bayside, via Newfound Harbor Channel, through the bight to the Atlantic Ocean as well as the coastal waters flowing towards Coupon Bight and Big Pine Channel on an incoming tide. |

All datasonde locations within Florida Keys Aquatic Preserves have very minimal freshwater input (i.e. rainfall only). These areas are tidal influenced, however tidal range is dependent on the land formations nearby.

Florida Keys Station Deployment Timeline:

| Station Code | Station Name | Location | Active Dates | Reason for Decommission | Notes |
| --- | --- | --- | --- | --- | --- |
| FKLK | Lignumvitae Key Aquatic Preserve (LKAP) | 24.90345, -80.69456 | 10/31/2021 - Present | N/A | EXO2 |
| FKCB | Coupon Bight Aquatic Preserve (CBAP) | 24.62045, -81.40563 | 10/31/2023 – Present | N/A | EXO2 |

6)  Data collection period –

FKLK

| Deployment Date/Time |  | Retrieval Date/Time |  | Notes |
| --- | --- | --- | --- | --- |
| 1/26/26 | 12:30 | 2/25/26 | 10:15 | * Failed Turbidity, DO CCV |
| 2/25/26 | 10:15 | 3/27/26 | 10:45 | * Failed Chlorophyll CCV |
| 3/27/26 | 10:45 | 4/20/26 | 9:00 |  |
| 4/20/26 | 9:00 | 5/22/26 | 9:30 |  |
| 5/22/26 | 9:30 | 6/15/26 | 11:30 | *Chl-a drift, changed to RFU, failed DO% and mg/L CCV |
| 6/15/26 | 11:30 | 7/24/26 | 9:15 | *Failed Chl-a, DO% and mg/L CCV |

FKCB

| Deployment Date/Time |  | Retrieval Date/Time |  | Notes |
| --- | --- | --- | --- | --- |
| 12/10/25 | 13:30 | 2/3/26 | 8:15 | * Failed Chlorophyll CCV, deployed 55 days due to weather |
| 2/3/26 | 8:15 | 3/9/26 | 12:30 | * Failed Chlorophyll CCV |
| 3/9/26 | 12:30 | 04/6/26 | 12:00 | * Failed DO CCV |
| 4/6/26 | 12:00 | 5/8/26 | 10:30 | *Changed Chl-a to RFU |
| 5/8/26 | 10:30 | 6/1/26 | 11:45 | *Failed DO% and mg/L CCV |
| 6/1/26 | 11:45 | 6/25/26 | 9:00 | *Failed Chl-a |
| 6/25/26 | 9:00 | 7/22/26 | 2:30 |  |

* All applicable parameter data flagged as suspect

** Catastrophic instrument and/or battery malfunction; all applicable data lost or rejected

Note: Deployment and retrieval times are based on when the datasonde began collecting data and when it was replaced with another datasonde. All time is reported as Eastern Standard Time.

NOTE: Specific conductance data is subject to occasional single ‘dips’ of reduced concentrations occurring anytime throughout a deployment. This decrease is most likely attributed to debris or live critters disrupting the signal being sent between the electrodes and the Conductivity/Temperature sensor during sample collection. These instances are coded as suspect when they do not affect other parameters. In cases where other related parameters are impacted, the Conductivity and its associated parameters (Salt, DO mg/L, and Depth) are rejected.

7)  Distribution –

The Principal Investigator (PI) retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the Florida Department of Environmental Protection, and the Aquatic Preserve site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.

Aquatic Preserve water quality data and metadata can be obtained from the PI and online at the Aquatic Preserve’s data portal home page www.floridaapdata.org.  Data are available in comma delimited format.

8)  Associated researchers and projects –

There are several agencies conducting water quality monitoring within the Florida Keys National Marine Sanctuary (FKNMS). The following are the projects that we are aware of though there may be more.

(This Project) Florida Department of Environmental Protection (DEP) – FKNMS/Florida Keys Aquatic Preserves (FKAP)

- Currently, two deployed datasonde stations monitor abiotic parameters at 15-minute intervals.

- In addition to the datasondes, quarterly water quality grab samples are taken at four locations within LKAP (two oceanside and two bayside) and three within CBAP (two in the bight and one oceanside). Field measurements include temperature, dissolved oxygen, pH, salinity, and conductivity. Laboratory measurements include chlorophyll-a, pheophytin-a, total phosphorous, nitrate-nitrite, total Kjeldahl nitrogen, and turbidity.

Florida International University (FIU)

- Long-term water quality sampling within FKNMS as part of the Southeast Environmental Research Center (SERC) Water Quality Monitoring Network.

  - http://serc.fiu.edu/wqmnetwork/

- FIU contract via the seagrass lab through Dr. Jim Fourqurean to assess FKNMS/FKAP collected samples of sediments and seagrass tissues.

  - https://seagrass.fiu.edu/

South Florida Water Management District (SFWMD)

- Long-term surface water quality monitoring in Florida Bay.

- Flow monitoring from Everglades National Park into Florida Bay.

- Data stored and accessible through DBHYDRO database.

  - http://my.sfwmd.gov/dbhydroplsql/show_dbkey_info.main_menu

Florida Department of Health (DOH)

- DOH Healthy Beaches program

- Sample monitoring for fecal indicator bacteria.

  - http://www.floridahealth.gov/environmental-health/beach-water-quality/index.html

II.  Physical Structure Descriptors

9)  Sonde and Sensor specifications –

YSI Exo2 Datasondes

| No. | Name | Model No. | Serial No. | Date Initiated | Service Date |
| --- | --- | --- | --- | --- | --- |
| 1 | Colonel Sonde-rs | 599502-01 | 21D104496 | 6/30/21 | 12/18/24 |
| 2 | Kiefer Sonde-rland | 599502-01 | 21D104497 | 10/12/21 | 5/12/25 |
| 3 | Sonde-ra Bullock | 599502-01 | 21D104498 | 10/12/21 | 4/30/25 |
| 4 | Samuel L. Jack-sonde | 599502-01 | 21D104499 | 1/5/22 | 2/10/25 |
| 5 | Pamela Ander-sonde | 599502-01 | 22E101548 | 10/11/22 | 3/14/25 |
| 6 | Sonde-Claude Van Damme | 599502-01 | 22E105877 | 10/11/22 | 4/30/25 |
| 8 | Mary Kate Ol-sonde | 599502-02 | 23F105038 | 10/31/2023 | 10/8/25 |
| 9 | Ashley Ol-sonde | 599502-02 | 23F105039 | 10/31/2023 | 11/24/25 |
| 10 | Dwane “The Rock” John-sonde | 599502-01 | 23B105131 | 11/14/2023 | 11/25/24 |
| 11 | Jim Par-sondes | 599502-01 | 23B105132 | 1/17/2024 | 2/17/25 |
| 12 | Owen Wil-sonde | 599502-01 | 23B105133 | 2/23/2024 | 2/21/25 |
| 13 | Adam Sondler | 599502-01 | 25A106828 | 3/24/2025 | N/A |

YSI EXO2 Datasonde Parameters and Acceptance Criteria:

- Parameter: Temperature

  - Units: Celsius (°C)

  - Sensor Type: 4-electrode nickel cell

  - Model#: YSI 599827

  - Range: -5 to 35 ºC, 35 to 50 ºC

  - Accuracy: -5 to 35 ºC: +/- 0.01 ºC2, 35 to 50 ºC: +/- 0.05 ºC2

  - Resolution: 0.001 ºC

  - Acceptance Criteria: ±0.3 °C of a NIST-traceable thermometer

- Parameter: Conductivity

  - Units: milli-Siemens per cm (mS/cm)

  - Sensor Type: 4-electrode nickel cell

  - Model#: YSI 599827

  - Range:0 to 200 mS/cm

  - Accuracy: 0 to 100: +/- 0.5% of reading or 0.001 mS/cm, w.i.g.; 100 to 200 +/- 1% of reading

  - Resolution: 0.0001 to 0.01 mS/cm

  - Acceptance Criteria: ±5%

- Parameter: Salinity

  - Units: practical salinity units (psu)/parts per thousand (ppt)

  - Model#: 599827

  - Sensor Type: Wiped probe; Calculated from conductivity and temperature

  - Range: 0 to 70 ppt

  - Accuracy: ±2% of the reading or 0.2 ppt, whichever is greater

  - Resolution: 0.01 psu

  - Acceptance Criteria: See Conductivity

- Parameter: Dissolved Oxygen, % air saturation

  - Units: percent air saturation (%)

  - Sensor Type: optical sensor

  - Model#: YSI 599100-01

  - Range: 0 to 500% air saturation

  - Accuracy: 0 to 200%: +/1 1% of reading or 1% saturation, w.i.g.; 200 to 500% +/- 5% of reading

  - Resolution: 0.1% air saturation

  - Acceptance Criteria: ±5%

- Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature, and salinity)

  - Units: milligrams/Liter (mg/L)

  - Sensor Type: Optical probe w/ mechanical cleaning

  - Model#: 599100-01

  - Range: 0 to 50 mg/L

  - Accuracy: 0-20 mg/L: ±0.1 mg/L or 1% of the reading, whichever is greater; 20-50 mg/L: ± 5% of the reading, relative to calibration gasses

  - Resolution: 0.01 mg/L

  - Acceptance Criteria: ±0.1 mg/L

- Parameter: Depth (non-vented, shallow)

  - Units: meters (or ft)

  - Sensor Type: integral, non-vented depth sensor

  - Range: 0 to 10 m (0 to 33 ft)

  - Accuracy: +/- 0.004 m (+/- 0.013 ft)

  - Resolution: 0.001 m (0.001 ft)

  - Acceptance Criteria:

- Parameter: pH

  - Units: pH units

  - Sensor Type: unguarded

  - Model#: YSI 599702

  - Range: 0 to 14 units

  - Accuracy: +/- 0.1 pH units within +/- 10ºC of calibration temp; +/- 0.2 pH units for entire temp range

  - Resolution: 0.01 units

  - Acceptance Criteria: ±0.1

- Parameter: Turbidity

  - Units: formazin nephelometric units (FNU)

  - Sensor Type: optical sensor

  - Model#: YSI 599101-01

  - Range: 0 to 4000 FNU

  - Accuracy: 0 to 999 FNU: 0.3 FNU or +/- 2% of reading, w.i.g.; 1000 to 4000 FNU: +/- 5% of reading

  - Resolution: 0 to 999 FNU=0.01 FNU; 1000 to 4000 FNU=0.1 FNU

  - Acceptance Criteria: whichever is greater ±0.3 FNU or ±5% of the standard value

- Parameter: Total Algae

  - Units: micrograms/Liter (mg/L)

  - Sensor Type: optical sensor

  - Model#: YSI 599103-01

  - Range: 0 to 400 µg/L chl

  - Accuracy: Linearity: r2  ≥ 0.999 for Rhodamine WT across full range

  - Resolution: 0.01 RFU or 0.01 µg/L chl

  - Acceptance Criteria: ±5%

Sensor Disclaimers:

Depth Qualifier:

YSI data sondes can be equipped with either vented or non-vented depth/level sensors.  Readings for both vented and non-vented sensors are automatically compensated for water density change due to variations in temperature and salinity; but for all non-vented depth measurements, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.02 cm for every 1 millibar change in atmospheric pressure and is eliminated for vented sensors because they are vented to the atmosphere throughout the deployment time interval.

Standard calibration protocol calls for all non-vented depth sensors to read 0 meters at a (local) barometric pressure of 1013.25 mb (760 mm/hg).  To achieve this, each site calibrates their depth sensor with a depth offset number, which is calculated using the actual atmospheric pressure at the time of calibration and the equation provided in the Aquatic Preserve calibration sheet or digital calibration log.  This offset procedure standardizes each depth calibration. If accurate atmospheric pressure data are available, non-vented sensor depth measurements can be corrected. The Principal Investigator should be contacted in order to obtain information regarding atmospheric pressure data availability.

Salinity Units Qualifier:

The EXO sondes report practical salinity units (psu). These units are essentially the same as parts per thousand (ppt) and for Aquatic Preserve purposes are understood to be equivalent, however psu is considered the more appropriate designation. Moving forward the Aquatic Preserve program will assign psu salinity units for all data regardless of the sonde type.

Turbidity Qualifier:

The EXO sondes use formazin nephelometric units (FNU). These units are essentially the same as nephelometric turbidity units (NTU) but indicate a difference in sensor methodology. For Aquatic Preserve purposes they will be considered equivalent.  Moving forward, the Aquatic Preserve program will use FNU/NTU as the designated units for all turbidity data regardless of the sonde type. If turbidity units and sensor methodology are of concern, please see the Sensor Specifications portion of the metadata.

Chlorophyll Fluorescence Disclaimer:

YSI chlorophyll sensors are designed to serve as a proxy for chlorophyll concentrations in the field for monitoring applications and complement traditional lab extraction methods; therefore, there are accuracy limitations associated with the data that are detailed in the YSI manual including interference from other fluorescent species, differences in calibration method, and effects of cell structure, particle size, organism type, temperature, and light on sensor measurements.

Conductivity/Depth/Dissolved Oxygen Disclaimer

Conductance data (specific conductivity) is subject to occasional single ‘dips’ of reduced concentrations occurring anytime throughout a deployment. This decrease is most likely attributed to debris or live organisms disrupting the signal being sent between the electrodes and the Conductivity/Temperature sensor during sample collection. These instances along with its associated parameters (salinity, dissolved oxygen mg/L, and depth) are rejected.

10)  Coded variable definitions –

| Sampling Station: | Sampling Site Code: | Station Code: |
| --- | --- | --- |
| Lignumvitae Aquatic Preserve | LKAP | FKLK |
| Coupon Bight Aquatic Preserve | CBAP | FKCB |

11)  QAQC flag definitions –

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is missing and above or below sensor range.  All remaining data are then flagged 0, passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3		Data Rejected due to QAQC

-2		Missing Data

-1		Optional SWMP Supported Parameter

0		Data Passed Initial QAQC Checks

1		Suspect Data

2		Open - reserved for later flag

3		Calculated data: non-vented depth/level sensor correction for changes in barometric pressure

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12)  QAQC code definitions –

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the deployment or YSI datasonde, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an * below) can be applied to the entire record in the F_Record column.

General Errors

GIC	No instrument deployed due to ice

GIM	Instrument malfunction

GIT	Instrument recording error; recovered telemetry data

GMC 	No instrument deployed due to maintenance/calibration

GNF	Deployment tube clogged / no flow

GOW	Out of water event

GPF	Power failure / low battery

GQR	Data rejected due to QA/QC checks

GSM	See metadata

Corrected Depth/Level Data Codes

GCC	Calculated with data that were corrected during QA/QC

GCM	Calculated value could not be determined due to missing data

GCR	Calculated value could not be determined due to rejected data

GCS	Calculated value suspect due to questionable data

GCU 	Calculated value could not be determined due to unavailable data

Sensor Errors

SBO	Blocked optic

SCF	Conductivity sensor failure

SCS		Chlorophyll spike

SDF	Depth port frozen

SDG	Suspect due to sensor diagnostics

SDO	DO suspect

SDP	DO membrane puncture

SIC		Incorrect calibration / contaminated standard

SNV	Negative value

SOW	Sensor out of water

SPC	Post calibration out of range

SQR	Data rejected due to QAQC checks

SSD	Sensor drift

SSM	Sensor malfunction

SSR		Sensor removed / not deployed

STF	Catastrophic temperature sensor failure

STS		Turbidity spike

SWM	Wiper malfunction / loss

Comments

CAB*	Algal bloom

CAF	Acceptable calibration/accuracy error of sensor

CAP	Depth sensor in water, affected by atmospheric pressure

CBF	Biofouling

CCU	Cause unknown

CDA*	DO hypoxia (<3 mg/L)

CDB*	Disturbed bottom

CDF	Data appear to fit conditions

CFK*	Fish kill

CIP	*	Surface ice present at sample station

CT*	Low tide

CMC*	In field maintenance/cleaning

CMD*	Mud in probe guard

CND	New deployment begins

CRE*	Significant rain event

CSM*	See metadata

CTS	Turbidity spike

CVT*	Possible vandalism/tampering

CWD*	Data collected at wrong depth

CWE*	Significant weather event

13)  Post deployment information –

CCV values in red did not pass the established acceptance criteria in Section 9 and data from the deployments were labeled as suspect.

| FKLK |  | Specific Conductivity (mS/cm) |  | pH |  | Temperature (°C) |  | Dissolved Oxygen (mg/L) |  | Turbidity (FNU) |  | Chlorophyll-a (µg/L) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sonde | Date | 50 | 70 or 100 | 7 | 10 | Sonde | Therm. | DO CCV | Std | 124 | 0 | 0 | Chl CCV | Std |
| 3 | 1/26/26 | 49.152 | 100.606 | 6.88 | 9.90 | 22.022 | 21.9 | 9.01 | 8.743 | 129.72 | 1.06 | 0.33 | 64.20 | 64.0 |
| 8 | 2/25/26 | 50.138 | 101.462 | 7.04 | 10.05 | 19.578 | 19.3 | 9.69 | 9.18 | 70.52 | -0.08 | -0.04 | 67.34 | 68.16 |
| 5 | 3/27/26 | 49.748 | 100.770 | 7.03 | 10.01 | 20.909 | 20.7 | 9.05 | 9.00 | 124.95 | 0.17 | -0.20 | 70.5 | 66.48 |
| 1 | 4/20/26 | 49.982 | 100.646 | 7.01 | 9.97 | 20.371 | 20.3 | 9.35 | 9.092 | 122.21 | .02 | 0.00 | 69.08 | 66.72 |
| 10 | 5/22/26 | 49.849 | 100.787 | 7.06 | 10.04 | 20.713 | 20.5 | 9.19 | 9.00 | 123.68 | 0.00 | 0.02 | 64.69 | 65.75 |
| 6 | 6/15/20 | 49.448 | 100.074 | 7.00 | 9.97 | 20.959 | 20.8 | 9.60 | 9.00 | 124.64 | .50 | 0.04 | 16.59 | 16.4 |

| FKCB |  | Specific Conductivity (mS/cm) |  | pH |  | Temperature (°C) |  | Dissolved Oxygen (mg/L) |  | Turbidity (FNU) |  | Chlorophyll-a (µg/L) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sonde | Date | 50 | 70 or 100 | 7 | 10 | Sonde | Therm | DO CCV | Std | 124 | 0 | 0 | Chl CCV | Std |
| 11 | 2/3/26 | 50.659 | 99.943 | 6.98 | 9.98 | 19.485 | 19.4 | 9.34 | 9.276 | 125.13 | 0.49 | 0.02 | 63.00 | 66.96 |
| 1 | 3/9/26 | 50.893 | 101.030 | 7.14 | 9.98 | 20.226 | 20.2 | 9.16 | 9.092 | 123.43 | 0.11 | 0.01 | 62.17 | 66.6 |
| 11 | 4/6/26 | 50.252 | 101.784 | 6.95 | 9.94 | 20.957 | 20.8 | 9.41 | 8.915 | 119.82 | 0.28 | 0.08 | 67.40 | 66.24 |
| 2 | 5/8/26 | 49.805 | 101.156 | 7.07 | 10.0 | 20.484 | 20.4 | 9.15 | 9.00 | 121.36 | -3.80 | 0.20 | 65.58 | 66.24 |
| 13 | 6/1/26 | 50.545 | 102.051 | 6.95 | 10.0 | 20.365 | 21.2 | 9.32 | 8.915 | 123.0 | 0.10 | -0.03 | 15.89 | 16.4 |
| 3 | 6/25/26 | 49.740 | 99.776 | 6.98 | 9.97 | 20.704 | 20.5 | 9.28 | 9.00 | 123.03 | 0.20 | 0.07 | 19.25 | 16.4 |

Temperature Continuous Calibration Verification (CCV) was not consistently performed, however unless marked otherwise all sensors passed temperature verification before next deployment.

14)  Other remarks/notes –

All data files are QAQC’d using the 2021 CDMO Excel macro. This macro automatically flags DO values less than 3 mg/L as Passed Initial QAQC Checks with the comment DO Hypoxia (<3 mg/L) <0> (CDA). The DO flag is based on Federal standards. During the primary QAQC session, any reading that falls outside a sensor’s range (see section II. Physical Structure descriptors) are rejected <-3>. Additionally, negative turbidity readings between (0) and (-2) have been automatically flagged as suspect data <1> (CAF) since the value is within the accuracy of the probe per the CDMO operations manual. All negative chlorophyll readings are labeled as <-4> since they fall outside the chlorophyll sensor range and are manually rejected during QAQC <-3> [SNV]. All the first data points for each new deployment are labeled (CND).

The parameters that are QC’d are depth (m), temperature (°C), salinity (ppt), specific conductance (mS), turbidity (FNU), pH (SU), DO (mg/L and %), and chlorophyll (µg/L). All other parameters have not been QC’d.

All data collected after the 30-days in each deployment is noted as “see metadata” (CSM).

Turbidity values above 124 FNU are flagged as suspect data due to being beyond the calibration range <-1> (CSM). Data users should exercise caution when interpreting turbidity data above 124 FNU.

Apparently anomalous data are rejected <-3> and labeled with cause unknown (CCU) if they are not within a well-defined peak (e.g., there is no incremental increase before data point or decrease after data point based on the neighboring readings) and the data point is more than 5 standard deviations away from the 24-hour average.

When swapping equipment, both datasondes were left in the water for one overlapping data point unless otherwise noted in section 6. However, only the data from the datasonde in the mount was retained beyond the raw data. The clipped data is colored grey in the electronic calibration field logs. If both sondes were outside of the mount, the data was clipped and the gap is labeled missing <-2> due to maintenance [GMC].

15) Acknowledgement:

The data included in this document was collected by the staff of the Florida Department of Environmental Protection. Any products derived from this data should clearly acknowledge this source (please use the attached logos). This recognition is important for ensuring that this long-term monitoring program continues to receive the necessary political and financial support.