## Sheet1
Code | Definition | Flags | Hints
General Errors |  |  | Used to document general problems with the deployment or common instrument related errors
GIC | No instrument deployed due to ice | -2 | Used for missing data as a result of winter/ice conditions
GIM | Instrument malfunction | -3 | Use for sonde failures
GIT | Instrument recording error; recovered telemetry data | 1, -3 | Use for telemetered data downloaded from the CDMO when raw sonde files are unavailable
GMC | No instrument deployed due to maintenance/calibration | -2 | Use for missing data in between deployments
GNF | Deployment tube clogged / no flow | 1, -3 | Use if water is not circulating through the deployment tube properly
GOW | Out of water event | -3 | Use if sonde is depoyed, but all probes were out of the water
GPF | Power failure / low battery | 1, -3 | Reject for power failures, some low battery situations may remain suspect
GQR | Data rejected due to QA/QC checks | -3 | Use for data reviewer determined should be rejected, generally applied to all parameter data
GSM | See metadata | Any | Points users to metdata to further explain a flag/code or provide additional information
Sensor errors |  |  | Used to document common sensor specific problems
SBO | Blocked optic | 1, -3 | Use when optics are blocked or partially blocked by wipers or foreign matter, generally reject
SCF | Conductivity sensor failure | -3 | Use for conductivity sensor failure, reject dependent parameters
SCS | Chlorophyll spike | 1, -3 | Use for outliers or discrete short-term events
SDF | Depth port frozen | 1, -3 | Use when a frozen depth port is indicated, generally reject
SDG | Suspect due to sensor diagnostics | 1 | Use for entire deployment when diagnostics (pH slope, conductivity cell constant, etc.) are out of range
SDO | DO suspect | 1 | Use when DO appears suspect but there is no evident cause
SDP | DO membrane puncture | 1, -3 | Use for rapid-pulse DO membrane punctures, generally reject, generally no longer applicable for optical DO probes
SIC | Incorrect calibration / contaminated standard | 1, -3 | Use for calibration issues, known calibration errors MAY be correctable (5 flag)
SNV | Negative value | -3 | Use for negative values that are not small acceptable turbidity or depth readings
SOW | Sensor out of water | -3 | Use (if sonde is deployed) for individual sensors that are out of the water and their dependent parameters
SPC | Post calibration out of range | 1, -3 | Reject if data appear significantly impacted, mark suspect if data appear minimally impcated, continue flagging/coding through end of deployment
SQR | Data rejected due to QAQC checks | -3 | Use for sensor specific data reviewer determined should be rejected (for issues not covered by other more specific codes like SPC or CBF)
SSD | Sensor drift | 1, -3 | Use for sensor drift where post-deployment checks are not out of range, apply to periods of data (not intermittently) and continue through end of deployment
SSM | Sensor malfunction | 1, -3 | Use for sensor specific malfunctions
SSR | Sensor removed / not deployed | -2, -3 | Use -3 flags if values are reported in the raw file
STF | Catastrophic temperature sensor failure | -3 | Reject all dependent parameters (except Turb with EXO sondes)
STS | Turbidity spike | 1, -3 | Use for outliers or discrete short-term events that warrant a suspect or rejected flag
SWM | Wiper malfunction / loss | 1, -3 | Use in conjunction with the suspect flag if data do not appear significantly impacted, reject if they are
Parameter comments |  |  | Used to futher describe conditions for the user - most are self explanatory
CAB* | Algal bloom | Any
CAF | Acceptable calibration/accuracy error of sensor | 1 | Use for small negative turbidity (macro does automatically) and depth readings
CAP | Depth sensor in water, affected by atmospheric pressure | 1 | Use with negative depth data that is positive after correction for barometric pressure
CBF | Biofouling | Any | Generally a 1 or -3 flag, but in some cases may be used with 0 if fouling seemed to be resolved. Apply to periods of data, not intermittently.
CCU | Cause unknown | Any
CDA* | DO hypoxia (<3 mg/L) | Any | Automatically applied by the macro, may be overwritten with higher priority codes if data will be rejected.
CDB* | Disturbed bottom | Any | Use in particular with turbidity when weather or human impacts distrub the bottom.
CDF | Data appear to fit conditions | Any | Used to indicate that an outlier or event fit field conditions
CFK* | Fish kill | Any | Use when a fish kill was present in the area
CIP* | Surface ice present at sample station | Any | Use to indicate the presence of ice
CLT* | Low tide | Any | Use to indicate low tide if impacts are visible in data (extreme low water, etc.)
CMC* | In field maintenance/cleaning | Any | 1 or -3 if data are impacted, 0 if this indicates the first reading after maintenance was performed
CMD* | Mud in probe guard | Any
CND | New deployment begins | Any | Applied automatically during data upload, if pre/post deployment data are moved later, apply with macro (NOTE: all data records must go through primary QC/upload)
CRE* | Significant rain event | Any | Use when data are impacted 'significantly' by a rain event that should be highlighted, apply to periods of data for parameters or entire records
CSM* | See metadata | Any | Points users to metdata to further explain a flag/code or provide additional information
CTS | Turbidity spike | 0 | Use for minor outliers or discrete events that don't warrant a suspect or rejected flag
CVT* | Possible vandalism/tampering | Any | Generally would warrant 1 or -3
CWD* | Data collected at wrong depth | Any | 1 for depth, may be 1 or 0 for other parameters, in extreme cases reject all but retain depth if other parameters are not rejected
CWE* | Significant weather event | Any | Use when data are impacted 'significantly' by a weather event that should be highlighted, apply to periods of data for parameters or entire records
Record Comments
* | Comment codes marked with an asterisk above may also be used in the F_Record column